Technical SOP

| Title | Hemoccult® (Fecal occult blood) |  |
| --- | --- | --- |
| Prepared by | Omar Dualeh | Date: 06/03/2009 |
| Owner | Ruben Layug | Date: 7/30/12 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	ANALYTICAL PRINCIPLE	3

3.	SPECIMEN REQUIREMENTS	3

4.	REAGENTS	4

5.	CALIBRATORS/STANDARDS	5

6.	QUALITY CONTROL	5

7.	EQUIPMENT and SUPPLIES	6

8.	PROCEDURE	7

9.	CALCULATIONS	7

10.	REPORTING RESULTS AND REPEAT CRITERIA	7

11.	EXPECTED VALUES	8

12.	CLINICAL SIGNIFICANCE	8

13.	PROCEDURE NOTES	9

14.	LIMITATIONS OF METHOD	9

15.	SAFETY	9

16.	RELATED DOCUMENTS	10

17.	REFERENCES	10

18.	REVISION HISTORY	10

19.	ADDENDA	11

# TEST INFORMATION

| Assay | Method/Instrument | Order Code | Local Code |
| --- | --- | --- | --- |
| Fecal occult blood | Hemoccult® test | SOB |  |

Synonyms/Abbreviations

Hemoccult

Department

Urinalysis

# ANALYTICAL PRINCIPLE

The Hemoccult® test is based on the oxidation of guaiac by hydrogen peroxide to a blue-colored compound. The heme portion of hemoglobin, if present in the fecal specimen, has peroxidase activity which catalyzes the oxidation of alpha guaiaconic acid

(active component of the guaiac paper) by hydrogen peroxide (active component of the developer) to form a highly conjugated blue quinone compound.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None. |
| Specimen Collection and/or Timing | - Collect fecal specimens from bowel movements on three different days. - Separate samples should be taken from two different sections of each fecal specimen. - Return the completed test card to the physician or laboratory within 14 days after the first sample collection. |
| Special Collection Procedures | - The Hemoccult® test requires only a small fecal specimen. - The specimen is applied to the guaiac paper of the Hemoccult® Slide as a THIN SMEAR using the Applicator Stick provided. - Hemoccult® Slides are best developed no sooner than 3 days after sample application. This allows any fruit and vegetable peroxidases present in the sample to degrade. - Slides containing samples may be stored for up to 14 days at room temperature (15 to 30°C) before developing. |
| Other | Patients should follow the PATIENT INSTRUCTIONS - At least 7 days prior to and continuing through the test period. Roughage in the diet can increase test accuracy by helping uncover “silent” lesions which bleed intermittently |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Specimen Inoculated Hemoccult ® card |  |
| -Other Acceptable | None |  |
| Collection Container | See section 3.1. |  |
| Volume | Specimen Inoculated Hemoccult ® card |  |
| - Optimum |  |  |
| - Minimum | Specimen Inoculated Hemoccult ® card |  |
| Transport Container and Temperature | Hemoccult® Slides (Test Cards). Transport to lab at Room Temp. |  |
| Stability & Storage Requirements |  |  |
|  | Room Temperature: | 14 days before developing |
|  | Refrigerated: | Unacceptable |
|  | Frozen: | Unacceptable |
| Timing Considerations | Hemoccult® Slides are best developed no sooner than 3 days after sample application however if immediate testing is required, wait 3 to 5 minutes before developing to allow adequate time for sample to penetrate the test paper. |  |
| Unacceptable Specimens & Actions to Take | Any other specimen other than fecal specimen. Recollect specimen. |  |
| Compromising Physical Characteristics | Protect slides from heat, light, and volatile chemicals. Refer to section 14.3. |  |
| Other Considerations | Because this test is visually read and requires color differentiation, it should not be interpreted by individuals with blue color deficiency (blindness). |  |

# REAGENTS

4.1	Reagent Summary

| Reagents | Supplier & Catalog Number |
| --- | --- |
| Hemoccult® Slides (Test Cards) and Developer | Beckman Coulter Product No. 60152 |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | Hemoccult® Slides (Test Cards) or Tape (in plastic dispenser) containing guaiac paper |
| --- | --- |
| Container | Single Slides card |
| Storage/ Stability | Slides will remain stable until the expiration dates which is printed on each slide |
| Preparation | None, ready to use. |

| Reagent | Hemoccult® Developer– a developing solution containing a stabilized mixture of less than 5.0% hydrogen peroxide and 75% denatured ethyl alcohol in an aqueous solution |
| --- | --- |
| Container | Single Slides card |
| Storage/ Stability | Developer will remain stable until the expiration dates which is printed on each bottle. |
| Preparation | None, ready to use. |

# CALIBRATORS/STANDARDS

N/A

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| Internal QC (The Performance Monitor areas) | Beckman Coulter Product # 60152 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | The Performance Monitor areas |
| --- | --- |
| Preparation | Apply one drop of Hemoccult® Developer between the Positive and negative Performance Monitor® areas. |
|  | • Read results within 10 seconds. If the slide and developer are functional, a blue color will appear in the positive Performance Monitor® area and no blue will appear in the negative Performance Monitor® area. |
|  | • Neither the intensity nor the shade of the blue from the Positive Performance Monitor® area should be used as a reference for the appearance of positive test results. |
|  | • Any blue originating from the positive Performance Monitor® area should be ignored when reading the sample test results. |
| Storage/Stability | Until expiration date printed on the label when stored at 15oC to 30oC |

6.3	Frequency

The Performance Monitor® areas must be developed on every slide.

6.4	Tolerance Limits

6.4.1	All quality control results are recorded on the Hemoccult work sheet.

6.4.2	Any quality control failures will be investigated using the Providence Hospital Lab Quality Control Procedure as a guide.

- The first step in the investigation will be repeating the test.

- Confirm the expiration date and the lot# of all controls

- Check the performance of a known positive control material.  If the proper result is not obtained, discard and retest with fresh product.

6.5	Review Patient Data

Review patient results for unusual, trends or distributions, such as an unusually high percentage of abnormal results.

6.6	Documentation

Refer to the local policies and procedures for QC documentation and for record retention requirements.

6.7	Quality Assurance Program

    - Refer to   polices and procedures for quality assurance activities applicable to this procedure in QUEST QA manual.

    - The laboratory participates in CAP proficiency testing and in other State and Local agency PT programs when applicable.

    - CAP Proficiency Challenge

    - All personnel performing the assay must complete annual competency and training process and results must be kept on file.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

N/A

7.2	Equipment

N/A

  - Supplies

See section 4.1

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Developing the Test |
| --- | --- |
| 1. | Slides are best developed no sooner than three days after sample application to allow for degradation of any fruit and vegetable peroxidases that may be present in the fecal sample. However, if immediate testing is required, wait 3 to 5 minutes before developing to allow adequate time for sample to penetrate the test paper. |
| 2. | Open back of slide and apply two drops of Hemoccult® Developer to guaiac paper directly over each smear. |
| 3. | Read results within 60 seconds. Any trace of blue on or at the edge of the smear is positive for occult blood |

| 8.2 | Developing the Performance Monitor® (Quality Control) |
| --- | --- |
| 1 | Apply one drop of Hemoccult® Developer between the Positive and Negative Performance Monitor® areas. |
| 2 | Read results within 10 seconds. If the slide and developer are functional, a blue color will appear in the positive Performance Monitor® area and no blue will appear in the negative Performance Monitor® area |
| 3 | Neither the intensity nor the shade of the blue from the Positive Performance Monitor® area should be used as a reference for the appearance of positive test results. |
| 4 | Any blue originating from the positive Performance Monitor® area should be ignored when reading the sample test results. |
| 5 | Neither the intensity nor the shade of the blue from the Positive Performance Monitor® area should be used as a reference for the appearance of positive test results. |

# CALCULATIONS

Not applicable

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

- If the slide and developer are functional, a blue color will appear in the positive  	performance monitor area and no blue color will appear in the negative performance 	monitor area.

- Neither the intensity nor the shade of the blue from the positive performance monitor 	should be used as a reference for the appearance of positive test results.

- Any blue color originating from the positive performance monitor area should be 	ignored when reading the patient sample test area.

- If the performance monitors fail to give appropriate results the patient test slide must 	not be used to interpret patient test results.  The slide must be discarded and a new 	sample obtained.  Contact the floor for new slides.  Corrective action measures 	must be taken and documented in the Corrective action log in Quality Control 	Sheet for Hemoccult Tests.

10.2	Rounding

N/A

  - Units of Measure

N/A

10.4	Clinically Reportable Range (CRR)

None established

  - Repeat Criteria

If there is no trace of blue on or at the edge of the smear repeat the test.

# EXPECTED VALUES

11.1	Reference Ranges

Negative

11.2	Priority 1 & 2 Limits

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

The Hemoccult® test is a rapid, convenient and qualitative method for detecting fecal occult blood which may be indicative of gastrointestinal disease. It is not a test for colorectal cancer or any other specific diseases.

The Hemoccult® test is recommended for professional use as a diagnostic aid during routine physical examinations, for hospital patients to monitor for gastrointestinal bleeding in patients with iron deficiency anemia or recuperating from surgery, peptic ulcer, ulcerative

Colitis and other conditions, and in screening programs for colorectal cancer when the Patient Instructions are closely followed. Serial fecal specimen analysis is recommended when screening asymptomatic patients.

The Hemoccult® test and other unmodified guaiac tests are not recommended for use with gastric specimens.

# PROCEDURE NOTES

- FDA Status: Approved

- Validated Test Modifications: None

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

Not applicable

14.2	Precision

N/A

14.3	Interfering Substances

• Red meat (beef, lamb and liver)

• Aspirin (greater than 325 mg/day) and other non-steroidal anti-inflammatory drugs 		   such as ibuprofen, indomethacin and naproxen

• Corticosteroids, phenylbutazone, reserpine, anticoagulants, antimetabolites, and 	       		   cancer chemotherapeutic drugs

• Alcohol in excess

• The application of antiseptic preparations containing iodine (povidone/iodine      		    mixture) Dietary iron supplements will not produce false-positive test results with 		    Hemoccult® tests.11 Acetaminophen is not expected to affect test results.14 			    Substances which can cause false-negative test results:15

• Ascorbic acid (vitamin C) in excess of 250 mg per day

• Excessive amounts of vitamin C enriched foods (citrus fruits and juices)

• Iron supplements which contain quantities of vitamin C in excess of 250 mg per day

  - Clinical Sensitivity/Specificity/Predictive Values

N/A

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- If you drive on company business (in a company car, a rental, or your own personal vehicle) you must comply with the requirements in the "Driver Safety Manual".  Please remember that driving requires your FULL attention and multi-tasking while driving (e.g., talking on a cell phone) is an unacceptable risk that you must not take.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

Current Package Insert

# REFERENCES

Hemoccult® tests Package insert Beckman Coulter, Inc. September 2009.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| .000 | 01/18/10 | 19 | Updated Addendum | O Dualeh | R.Levy |
| .001 | 03/02/11 | Page 1 | Changed Owner | A. Foster | R. Levy |
|  |  |  | Removed effective date from laboratory approval |  |  |
|  |  | 3.1 | Editted specimen collection notation |  |  |
|  |  |  | Added other considerations to patient preparation |  |  |
|  |  | 8.0 | Remove Package insert from procedure notes |  |  |
|  |  | 16.0 | Added Package insert |  |  |
|  |  | 19.0 | Removed package insert from addendum |  |  |
| .002 | 7/30/12 | 1 | Changed owner. Changed annual review to review. | R. Layug | R. Levy |
|  |  | 17 | Updated package insert version to September 2009. |  |  |
|  |  | 19 | Updated Specimen/Quality Control Sheet |  |  |

# ADDENDA

| Addendum | Title |
| --- | --- |
| 1 | Specimen / Quality Control Sheet |