#### RAPID HIV-1/2 TEST (Oraquick ADVANCE)

| ##### POLICY & METHODOLOGY | - This HIV Rapid test is performed manually on Labor and Delivery patients, who have no record of previous HIV test. - The HIV Rapid Test will be used in Occupational Health department when needle stick injuries occur. [Note: Rapid HIV ½ Antibody test should be performed on source patients]. - OraQuick® ADVANCE Rapid HIV-1/2 Antibody Test is a single-use, CLIA-Waived, Qualitative Immunoassay to detect antibodies to Human Immunodeficiency Virus Type 1 [HIV-1] and Type 2 [HIV-2]. - This HIV Rapid test is not recommended for persons under 12 years of age. For this age group (<12 years old), the recommended screening test is the HIV ½ EIA. |
| --- | --- |
| ##### QUALITY ASSURANCE | - The Quality Assurance Plan for Rapid HIV testing is designed to: - Ensure accurate and reliable test results - Monitor the performance and effectiveness of the test system - Assure competency of testing personnel - Identify and correct problems associated with the test methodology - The Quality Plan assures the integrity of the pre-analytic, analytic, and post-analytic phases of the testing process by the following means: - Procedures are written and available to testing personnel describing all aspects of the testing process. - Integrity of test kits if verified by performing quality control on each new lot, each new shipment, and each day of patient testing thereafter. - The ability to perform a test properly is assured by appropriate training when a new test system is introduced, and annual skills evaluations thereafter. - Proficiency testing samples are distributed at scheduled times, and serve to assess operator performance. - Temperature of reagent storage and testing area is monitored and recorded to assure that testing takes place within the range specified by the manufacturer. - Quality Control and temperature records are reviewed monthly to verify that the test is being performed per protocol. Quality Improvement reports are generated to investigate and resolve problems associated with the performance of any lab test. Problems may be pre-analytic, analytic, or post-analytic. |

Continued on next page

| SPECIMEN AND STORAGE REQUIREMENTS | - Whole blood [with EDTA anti-coagulant] specimen collected by venipuncture is the required specimen for Rapid HIV 1/2 Test. - Simultaneously, another sample, SST serum tube is collected and accessioned for HIV 1/2 EIA - Testing should be performed only on the original (“virgin” or unopened) tube to avoid potential cross-contamination. Testing should not be performed on aliquots or tubes that have been previously entered. - If the specimens cannot be tested STAT or at the time of collection, the whole blood may be stored at refrigerated temperature [2-8ºC] for up to 30 hours. DO NOT freeze EDTA sample - Do not perform the HIV Rapid test if the collection time has exceeded the 30-hour limit. |
| --- | --- |

| ##### MATERIALS AND SUPPLIES | - Unused, unopened OraQuick® ADVANCE Rapid HIV-1/2 kits may be stored in the refrigerator, or at room temperature [2° - 27°]. - Do not open pouch until ready to perform test. - If stored refrigerated, bring to room temperature [15° - 37°] before opening. - The kit consists of: - Divided pouches containing a test device, absorbent packet and developer solution. - Reusable test stands - Specimen collection loops. |
| --- | --- |

| ##### QUALITY CONTROL [Q.C.] KIT & STORAGE | - OraQuick® ADVANCE Rapid HIV-1/2 Antibody Test Kit Controls - Are ordered separately - Must be stored at refrigerator temperature [2° - 8°C] - The QC kit consists of: - HIV-1 Positive Control, 1 vial, black cap - HIV-2 Positive Control, 1 vial, red cap - Negative control, 1 vial, white cap |
| --- | --- |

| ##### Other Materials required but not provided | - Timer - Clean, disposable, absorbent workspace cover - Biohazard waste container to dispose of used materials - Sterile gauze pads |
| --- | --- |

Continued on next page

| ##### SAFETY & HANDLING PRECAUTIONS | Follow the safety and handling precautions listed below when performing HIV Rapid Test. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Observe Standard precautions for handling infectious agents when performing this test. Handle blood specimens and materials contacting blood specimens as if capable of transmitting infectious agents. |
| 2. | Use all Specimen Collection Loops, Test Devices, and Developer Solution Vials only once and dispose of properly (see Safety Precautions). Do not reuse any of these test components |
| 3. | Do not use the test beyond the expiration date printed on the Divided Pouch. Always check expiration date prior to testing. |
| 4. | Do not interchange Test Devices and Developer Solution Vials from kits with different lot numbers. |
| 5. | Insert the Test Device into the Developer Solution Vial within 60 minutes after introducing the venipuncture whole blood sample, to ensure accurate test results. |

| # QUALITY CONTROL | External quality controls are to be tested as follows: - Whenever a new test kit lot # is received. - Whenever a new shipment of test kits is received. - Each day of patient testing. - Perform in the same manner as a patient test. Built in procedural control - A line at the Control (“C”) area of the result window indicates that the specimen was added, and has migrated appropriately through the test device. - The Control line will appear on all valid tests (both positive and negative) |  |
| --- | --- | --- |
| ##### SET UP YOUR WORKSPACE | - If kits are stored at refrigerated temperature, gather the materials you will need. - Allow the test kit and quality control kit to come to operating temperature [15–37C; 59–99F] before use. - Refer to the External Quality Control section above to determine when the Kit Controls should be run. - Cover your workspace with a clean, disposable, absorbent workspace cover. - Set an OraQuick ADVANCE Reusable Test Stand [“Stand”] up on your workspace cover. Use only the stand provided. |  |

Continued on next page

| ##### TEST PROCEDURE |  |
| --- | --- |

| STEP | ACTION |
| --- | --- |
| 1. | Open the OraQuick® ADVANCE Divided Pouch - Remove the Developer Solution Vial - Label with patient name. - Remove the cap by gently rocking back and forth and place the vial into a slot in the stand Note: DO NOT cover the two holes in the back of the Device with labels or other materials. Doing so may cause an Invalid result |
| 2. | Add the whole blood EDTA specimen as follows: - Prior to sampling, mix the tube gently to ensure a homogeneous sample. - Remove an aliquot of blood using an unused Specimen Collection Loop (if the loop is dropped or comes in contact with any other surface, discard it in a biohazard waste container. Obtain a new loop.) - Insert the blood-filled end of the loop into the developer vial and stir. - Discard loop in a biohazard container, and verify that the solution is pink (if the solution is not pink, start the test over using a new pouch). |
| 3. | Remove the Test Device from the pouch(do not touch pad): - Use only if the absorbent packet is present. - With the Result Window facing forward, insert the flat pad of the device into the vial. - The flat pad must touch the bottom of the vial. |
| 4. | Set the timer for 20 minutes - Do not remove the device from the vial while the test is running - Read the results between 20 and 40 minutes in a fully lighted area. |

Continued on next page

| ##### TEST RESULTING AND INTER-PRETATION | Refer to table below for test resulting and interpretation. |
| --- | --- |

| Step | Action |  |  |  |
| --- | --- | --- | --- | --- |
| 1 | Refer to the Result Window on the Test Device. |  |  |  |
|  |  | If the Result Window shows…. | Then the result is… |  |
|  |  | a reddish-purple line appears next to the triangle labeled “C”, and NO line appears next to the triangle labeled “T”… | Non-Reactive |  |
|  |  | A reddish-purple line appears next to the triangle labeled “C” and a reddish-purple line appears next to the triangle labeled “T”. [One of these lines may be darker than the other.] | Reactive |  |
|  |  | no reddish-purple line appears next to the triangle labeled “C”, or | Invalid |  |
|  |  | a red background in the Result Window makes it difficult to read the result after 20 minutes, or | Invalid |  |
|  |  | if any of the lines are NOT inside the “C” or “T” triangle areas | Invalid |  |

Continued on next page

| 2 | Test result is interpreted as follows: |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | If the result is… | Then it is interpreted as… |  |
|  |  | Non-Reactive [A Non-Reactive test result means that HIV-1 and HIV-2 antibodies were not detected in the specimen.]….. | Non-reactive for HIV-1 and HIV-2 antibodies. Medical Center Laboratory - reports Rapid HIV as non-reactive, and call L&D with the result. - retains the Rapid HIV 1/2 [whole blood EDTA] specimen(refrigerated). |  |
|  |  | A Reactive test result means that HIV-1 and/or HIV-2 antibodies have been detected in the specimen | Preliminary Positive for HIV-1 and/or HIV-2 antibodies. Medical Center Laboratory - reports Rapid HIV as Reactive, and call L&D with the result. - gives the SST7 tube to the Send Out Dept. to accession for Procedure code 9855400 [HIV-1 Western Blot] [Laboratory orderable only], and send the labeled Western Blot tube to the Regional Reference Laboratories. - retains the Rapid HIV 1/2 [whole blood EDTA] specimen(refrigerated). |  |
|  |  | An Invalid test result means that there was a problem running the test, either related to the specimen or to the Test Device | - Result cannot be interpreted. - Repeat the test with a new Divided Pouch and a new Whole blood sample. - Reject test using “ITRE” [technical error] - Contact Customer Service at OraSure Technologies. Inc., (1-800-672-7873) if you are unable to get a valid test result upon repeat testing. |  |

Continued on next page

| ##### TEST REPORTING | Use message codes below when reporting results for HIV Rapid test: |
| --- | --- |

| Step | Action |  |  |  |
| --- | --- | --- | --- | --- |
| 1. | Report the result as NHIV, PHIV or ITRE. |  |  |  |
|  |  | IF | THEN |  |
|  |  | NON- REACTIVE | Enter “NHIV” in LMS which translates into “Negative for HIV-1 and HIV-2 antibodies. |  |
|  |  | REACTIVE | Enter “PHIV” in LMS which translates into “Preliminary positive for HIV-1 and/or HIV-2 antibodies. Additional sample will be sent to the Regional Reference Laboratories for confirmation and final testing. |  |
|  |  | INVALID: Result cannot be interpreted. Patient to be redrawn. | Enter “ITRE” in LMS which translates into “Instrument/Technical-related error, test not performed”. |  |
| 2. | Write results in manual log in the urinalysis department. |  |  |  |

| ##### REFERENCE | Package insert of OraQuick ADVANCE Rapid HIV-1/2 Antibody Test [WORD file] Note: for current version refer to: http://www.orasure.com/products/ |
| --- | --- |