| 111 South Grant Ave. Columbus, OH 43215 | OHIOHEALTH LABORATORY SERVICES STANDARD OPERATING PROCEDURE |  |
| --- | --- | --- |
|  | TITLE: TCA SYVA RAPID TEST DAU ONE STEP |  |
|  | NUMBER GMC-CHM-MN-01 | EFFECTIVE: August 12, 2013 |

Intended Use:   To provide a qualitative preliminary analytical detection of Tricyclic Antidepressants in human urine.  The test detects nortriptyline at the cutoff of 1000ng/ml.

Summary:  Tricyclic Antidepressants are a type of prescription drug intended for clinically depressed patients.  These drugs are frequently abused and are a leading cause of drug overdose death in the United States.  Abuse of TCA’s may lead to coma respiratory depression, convulsions, blood pressure deviations, hyperpraxia and severe cardiac conditions.

There are two broad chemical classes of TCA’s.  The tertiary amines (amitriptyline, imipramine, trimipramine, and doxepin) boost serotonin levels and are prescribed for insomnia, irritability, and over stimulation.  The secondary amines (nortriptyline, desipramine and protriptyline) enhance norepinephrine levels and are prescribed for opposite types of symptoms, such as excessive fatigue, withdrawal and inertness.  TCA’s are excreted in the urine, mostly as metabolites, for up to ten days. 2, 3

Test Principle:  The Syva® RapidTest d.a.u.® TCA test employs one-step, solid-phase immunoassay technology to discretely detect the presence of tricyclic antidepressants in human urine. The test is based on the principle of the highly specific immunochemical reactions between antigens and antibodies, which are used for the analysis of specific substances in biological fluids. The test relies on competition for binding to the antibodies between drug conjugates and drugs that may be present in the urine sample.

In the test procedure, a sample of human urine is placed in the Sample (S) well of the device, and the sample is allowed to migrate upward. If present in the urine sample, the drug competes with the drug-dye conjugate for limited antibodies immobilized on the membrane. When a sufficient amount of drug is present, it will saturate the antibodies on the membrane. A drug positive sample will not generate a line at the TCA position in the result window, whereas a drug negative sample will generate a line at the TCA position.

In addition to the line that may appear at the TCA position in the result window, a Control line must appear at the Control (C) validation position in the result window to confirm the viability of the test. The Control line should always be seen if the test is conducted properly. The Control line contains immobilized polyclonal anti-mouse IgG; therefore, it will capture antibody-dye conjugates that pass the region, causing a colored line to show at the C position in the result window. This works as a procedural control, confirming that proper sample volume was used and the reagent system worked. If insufficient sample volume is used, there may not be a Control line, indicating the test is invalid.

Materials Provided

Each Syva® RapidTest d.a.u.® TCA (6A119UL) contains all the reagents necessary to perform one test.

- 1 Syva® RapidTest d.a.u.® TCA device. Each test device has (1) a membrane strip coated with sheep polyclonal antibody to nortriptyline and a control band coated with sheep antibody to mouse IgG, and (2) a dye pad containing colloidal gold coated with mouse IgG and a conjugate of nortriptyline in a protein matrix.

- 1 disposable sample dropper.

- Instructions for use.

Precautions

- For in vitro diagnostic use.

- Avoid cross contamination of urine samples by using a new urine specimen container and dropper for each urine sample.

- This device does not contain any human source material.

- Urine specimens are potentially infectious. Proper handling and disposal methods should be established according to good laboratory practice.

- The Syva® RapidTest d.a.u.® TCA device should remain in its original, sealed pouch until ready for use. Do not use the test if the pouch is damaged or improperly sealed.

- Do not use the test kit after the expiration date.

- This device contains nonsterile bovine serum albumin.

Storage and Stability

The Syva® RapidTest d.a.u.® TCA kit should be stored at 2–30°C (35–86°F) in the original, sealed pouch. The expiration dating was established under these storage conditions.

Specimen Collection and Preparation

Approximately 110 μL of urine sample is required for each test. Fresh urine specimens do not require any special handling or pretreatment. Specimens should be collected in a clean glass or plastic container. If testing will not be performed immediately, specimens should be refrigerated up to 48 hours (2–8°C) or frozen. Specimens should be brought to room temperature before testing.  Specimens containing a large amount of particulate matter may give inconsistent test results.  These specimens should be clarified before testing by centrifuging or allowing to settle.

Test Procedure

The test procedure consists of adding the urine sample to the Sample (S) well of the device and watching for the appearance of colored lines in the result window.

- For each test, open one pouch and label the device with the patient identification (ID).

- Holding the dropper vertically, dispense three drops (110 μL) of urine sample into the Sample (S) well.

- Read the result after five minutes, but within 10 minutes of sample addition.

- Log Patient results on Patient Result Log (Figure 2)

Interpretation of results

Negative:  the appearance of a reddish-purple Control ( C ) line and a line at the TCA position indicates a negative test result.  This means that no drug concentration above the cutoff level has been detected.

- Any faint line at the TCA position of the result window, visible within 10 minutes, should be interpreted as a negative result.

- A negative test result does not indicate the absence of drug in the sample:  In  qualitative terms, it only indicates that the sample does not contain drug above the cutoff level

NOTE:  The color intensities of the control line and the line at the TCA position may not be equal.

Positive: the appearance of a reddish-purple Control (C) line and no line at the TCA position in the result window indicates the test result is positive.  In qualitative terms, this means that the specimen contains the drug at a concentration above the cutoff level.

- A positive test result does not indicate the level of intoxication or the specific    concentration of drug in the sample.

Invalid:  A Control (C) line should always appear.  The test is invalid if no control line forms at the C position.  The test must be repeated with a new device.

Limitations:

- The test is designed for use with unadulterated human urine ONLY

- Other substances and/or factors not listed, such as technical or procedural errors, may interfere with the test and cause false results.

- Adulterants, such as bleach and/or alum, in urine specimens may produce erroneous results regardless of the method of analysis.  If adulteration is suspected, the test should be repeated using a new sample.

- The test result must be read within 5-10 minutes of sample application. The test result read after 10 minutes may not be consistent with the original reading obtained within the five to 10 minute reading period.

- Certain plastics may adsorb some drugs.

Internal Controls: Each Syva® RapidTest d.a.u.® TCA test device has built-in controls. The Control (C) line is an internal positive procedural control. A distinct reddish-purple Control line should always appear at the C position in the result window, if the test procedure is performed properly; an adequate sample volume is used; the sample and reagent are wicking on the membrane; and the test reagents at the control line and the conjugate-color indicator are reactive. In addition, if the test has been performed correctly and the device is working properly, the background in the result window will become clear and provide a distinct result. This may be considered an internal negative procedural control.

The positive and negative procedural controls contained in each Syva® RapidTest d.a.u.® TCA test device satisfy the requirements of testing a positive control and a negative control on a daily basis. If the Control line does not appear at the control position in the result window, the test is invalid and a new test should be performed. If the problem persists, contact Syva Company for technical assistance.

External Controls: External controls should also be used to ensure that the reagents and assay procedures are performing properly. Controls should be tested at regular intervals, as part of a good laboratory testing practice.

Use Biorad toxi control #454 as a negative control and Biorad toxi control #455 as a positive control once per lot/shipment.  In addition, a previous positive and previous negative patient must by run on a new lot of reagent

Log results on TCA control log (Figure 1)

Expected Values

Syva® RapidTest d.a.u.® TCA is a qualitative assay. The amount of tricyclic antidepressants present in human urine cannot be estimated by the test. The results distinguish drug-positive from drug-negative samples. Positive results indicate the samples contain tricyclic antidepressants at a concentration giving a response above the cutoff level.

Performance Characteristics

The Syva® RapidTest d.a.u.® TCA test detects a cutoff level of 1000 ng/mL nortriptyline in human urine.

Accuracy

Two hundred and three (203) samples were tested at a cutoff value of 1000 ng/mL nortriptyline.  The samples were analyzed using the Syva® RapidTest d.a.u.® TCA test and a commercially available immunoassay (Triage®), and the results were compared. Complete agreement was observed for 99% of the samples. Results are shown in Table 1.

Table 1 — Accuracy: Syva® RapidTest d.a.u.® TCA Test vs. Triage® Assays

Triage®                                       Positive        Negative                     TOTAL

Syva® RapidTest Positive         103                2                                 105

d.a.u.® TCA          Negative       0                    98                               98

TOTAL        103               100                              203

In an additional study, 20 blind samples were prepared by spiking various concentrations of nortriptyline into drug-negative specimens. Two operators independently tested the specimens using the Syva® RapidTest d.a.u.® TCA test. Results obtained by the operators were compared and showed complete agreement.

Precision

Syva® RapidTest d.a.u.® TCA test was determined by testing serially diluted

standard drug solutions. Approximately 95% of the samples containing nortriptyline

concentrations 25% greater than the cutoff level consistently showed positive results.

Cutoff precision was challenged by testing more than 40 samples with drug concentrations ± 25% of the cutoff level. Results obtained using the Syva® RapidTest d.a.u.® TCA test consistently agreed with expected test results.

Reproducibility

Performance of the Syva® RapidTest d.a.u.® TCA test was examined at three different sites.  Reproducibility was determined by testing 15 blind controls, consisting of five negative samples, five moderately positive samples (i.e., a concentration 1.5 times the cutoff level), and five strongly positive samples (i.e., a concentration 3 times the cutoff level). Results obtained at these three sites were compared and showed 100% agreement.

Specificity

Specificity was determined by using the Syva ® RapidTest d.a.u.® TCA test to analyze

drug-negative urine specimens spiked with various drugs and drug metabolites.

Table 2 lists the compounds detected by the panel and the concentrations at which the sample produce a positive result.

Table 2 — Specificity

Compound                                            Concentration(ng/mL)

Amitriptyline                                                         1000

Chlorpromazine				    	 85000

Clomipramine                                                        7500

Cyclobenzaprine                                                    1500

Desipramine                                                           1000

Diphenhydramine                                                50000

Dothiepin                                                                 125

Doxepin                                                                 1000

Imipramine                                                              850

Norclomipramine                                                 50000

Nordoxepin                                                            1000

Nortriptyline                                                          1000

Perphenazine                                                        40000

Promazine                                                            10000

Protriptyline                                                             400

Trimipramine                                                         1500

Table 3 lists the compounds, tested at a concentration of 100 μg/mL, that showed no cross reactivity with the Syva® RapidTest d.a.u.® TCA test.

Table3 — Specificity---Non Cross-Reacting compounds Compound

4-Acetamidophenol			(-) Cotinine

Acetophenetidin (Phenacetin)		Creatinine

N-Acetylprocainamide			Cyclizine

Acetylsalicylic acid			Deoxycorticosterone

Aminopyrine				Dextromethorphan

Amobarbital				Diazepam

Amoxapine				Diclofenac

Amoxicillin				Diethylpropion

D,L-Amphetamine			Diflunisal

L-Amphetamine			Digoxin

Apomorphine				Domperidone

Aspartame				Doxylamine

Atropine				Ecgonine

Benzilic acid				Ecgonine methylester

Benzoic acid				(+) Ephedrine

Benzoylecgonine			(+/-) Ephedrine

Benzphetamine				(-) Ephedrine

Butabarbital				(-) Y Ephedrine

Cannabidiol				Erythromycin

Cannabinol				B-estradiol

Chloralhydrate				Estrone-3-sulfate

Chloramphenicol			Ethyl-p-aminobenzoate

Chlordiazepoxide			Fenoprofen

Chloroquine				Furosemide			Iproniazid

Chlorothiazide				Gentisic acid			(–) Isoproterenol

Cholesterol				Glutethimide			Isoxsuprine

Clonidine 				Guaifenesin			Ketamine

Cocaine				Hippuric acid			Ketoprofen

Cortisone				Hydralazine			Labetalol

Hydrochlorothiazide			3-Hydroxytyramine		Lidocaine

Hydrocortisone				Ibuprofen

o-Hydroxyhippuric acid			Iproniazid

Loperamide								Phenylpropanolamine

Loxapine succinate							Prednisolone

Meprobamate								Prednisone

Promethazine

D,L-Propranolol

Methadone								Propiomazine

p-Hydroxymethamphetamine						D-Propoxyphene

Methaqualone								D-Pseudoephedrine

Methoxyphenamine							Quinidine

(±) 3,4-Methylenedioxyamphetamine					Quinine

(±) 3,4-Methylenedioxymethamphetamine				Ranitidine

Methylphenidate							Salicylic acid

Methyprylon								Secobarbital

Nalidixic acid								Serotonin

Naltrexone								Sulfamethazine

Naproxen								Sulindac

Niacinamide								Temazepam

Nifedipine								Tetracycline

Norethindrone								∆8-THC

Noroxymorphone							∆9-THC

D-Norpropoxyphene						11-nor-∆9-THC-9-COOH

(–) Norpseudoephedrine						tetrahydrocortisone

Noscapine								Tetrahydrozoline

Nylidrin								Thiamine

D,L-Octopamine						Thienylcyclohexylpiperidine

Oxalic acid								Thioidazine

Oxazepam								D,L-Thyroxine

Oxolinic Acid								Tolbutamide

Oxymetazoline								Triamterene

Papaverine								D,L-Tryptophan

Penicillin-G			Phentermine				Tyramine

Pentazocine								D,L-Tyrosine

Pentobarbital			Phenytoin				Uric Acid

Phencyclidine			L-Phenylephrine			Verapamil

Phendimetrazine				Zomepirac

Phenelzine                              β-Phenylethylamine

References

- Hawks RL, Chiang CN, eds. Urine Testing for Drugs of Abuse, Rockville, MD; Substance Abuse and Mental Health Services Administration (SAMHSA), NIDA Research Monograph 73; 1986.

- Bickel MH, Poisoning by Tricyclic Antidepressant Drugs, Int. J. Clinical Pharmacol. 11(2): 145–176; 1975.

- Baselt RC. Disposition of Toxic Drugs and Chemicals in Man. 2nd ed., Davis, CA: Biomedical Publ.: 1982.

- DOCUMENT REVIEW and APPROVAL TO IMPLEMENT

Prepared by			______                                                Date _________

Manager/Director		______                                                Date __________

Site Medical Director	             ______                                                Date _________ □ NA

CLIA Lab Director		______                                                Date __________

- Review

| Date | Signature | Title |
| --- | --- | --- |

Annual Reviews

Document History

Originated: 	October 2008

Revised:  	Reformatted 10/09 and 8/13

Retired:

Electronic Copy:

/tmp/in/input.doc