Name:  _______________________________				            Date:  _____________________

Graded by: _____________________________                           Passing Score 80% (12/15): _____________________

I.  i-STAT - True (T) or False (F)

_______1.	If QC test falls outside expected range when repeated and procedural error has been ruled out, the 10 consecutive day QC validation must be performed.

_______2.	cTnI & CHEM8 Cartridges stored at room temp. for >14 days can be used for testing.

_______3.	The External Electronic Simulator is run monthly & if the Internal Simulator failed.

_______4.	Upload the i-STAT meter after patient testing to transmit results to Health Connect.

_______5. 	All i-STAT cartridges must be at room temp. for 1 hour/box or 5 minutes/cartridge prior to testing.

_______6. 	Cartridges may be returned to the refrigerator once they have been at room temperature.

_______7. 	Liquid QC testing is performed as needed.

_______8. 	Submit “POCT Result Correction Form” for any error correction problems with i-STAT testing.

_______9.  	All cartridges taken out of the refrigerator should be dated with the new expiration date.

______10.	All failed QC should be documented on the Corrective Action Log.

- Select the comment code(s) that you can use to REJECT your iSTAT test or result.

  - 0

  - None. Do not enter any code; when you dock the meter, it will automatically reject the test.

  - 1

  - ___

- How much blood sample do you need to collect in a dark green top tube to use for testing?

  - 2 ml minimum

  - You only need 2 drops

  - The green top tube must be at least half full.

  - Preferably a full green top tube or must be at least ¾ full

- Select the critical iSTAT result(s) among the following:

  - Potassium of 6.2mmol/L

  - Na of 160 mEq/L or Na of 118 mEq/L

  - Hct of 15% and Hgb of  5g/dL

  - Lactate of 2.4 mEq/L

- Level 3 QC was out of range; per policy you may repeat it only once. If the repeat fails what would be your next action?

  - Send patient samples to the lab for testing until QC problem is resolved.

  - Disregard QC failure and run your patients anyway.

  - Continue running patient samples as long as QC level 1 passed.

  - Document all QC failure in the out of limits log in the QC binder and notify the Lab.

- What do you do if you get an iSTAT result of Glucose = 39 mg/dL?

  - Verify the integrity of your blood sample (time of draw, sample ID, volume of blood collected).

  - Notify the Provider immediately.

  - Document your Critical Read Back Notification in the Critical Notification Flowsheet in KPHC.

  - You don’t need to document anything because you already did for the 1st encounter.