Sample Template for:

Comprehensive Policies, Processes and Procedures Manual

For use with the ACCU-CHEK® Inform II Glucose Monitoring System

|  | Disclaimer |
| --- | --- |

The content of this document is intended only as a guide to assist you in developing a comprehensive policies, processes and procedures manual.  This is an optional auxiliary tool.  At the time of publication, it reflects the then-current reagent package inserts and instrument operator’s manual for the product.  Additional ACCU-CHEK Inform II system usage information, which resides in bulletins, updated manuals/package inserts or any other valid labeling source, may not be present.

Roche Diagnostics makes no representation or warranty that by utilizing the information contained in these procedures, a facility will be in compliance with all applicable laws or regulations, nor is this guide intended to be used for purposes of ensuring compliance. Compliance with applicable laws or regulations remains solely the responsibility of the individual facility.

The updating, review and approval of procedures for all laboratory tests, as well as supplementing them with any other information necessary to perform a valid test, is the responsibility of the program director or his/her designee. It is his/her responsibility to revise the procedure used in the facility when procedure changes are announced by Roche Diagnostics, or when there is a change in laws, regulations or their interpretations. Roche Diagnostics hereby disclaims all liability for any facility's compliance with any laws or regulations.

If you have any questions regarding the content of this document, please contact your local representative or call the ACCU-CHEK Customer Care service center at 1-800-440-3638.

|  | Introduction |
| --- | --- |

Roche Diagnostics is pleased to offer this tool to assist you in developing your comprehensive policies, processes and procedures manual for the use of the ACCU-CHEK Inform II system in your facility.  This guide contains sample policies, process tables and procedures for use with the ACCU-CHEK Inform II system and is intended only as a guide and example.  It is important that you edit and make updates as may be required by your internal policies and changes in policies or related laws or regulations or their interpretation.  This is a working document that generally conforms to CLSI document GP02-A5 format for policies, processes and procedures.  All text in this template is editable to reflect information specific to your facility and preferences.  You will find blank spaces___ and <gray italicized text> throughout this document.  These are provided to alert you to areas that you will commonly want to customize and to provide examples of text for you to consider.  You will need to delete or adjust the font of the <gray italicized text> after you have customized your entries.

ACCU-CHEK Inform II Glucose Monitoring System

Comprehensive Policies, Processes and Procedures Manual

(Sample Template)

|  | Table of Contents |  |
| --- | --- | --- |
| Comprehensive Policies, Processes and Procedures Manual Approval |  |  |
| ACCU-CHEK Inform II System |  |  |
| Extent of Use Policy |  |  |
| Principle of the Method |  |  |
| Limitations of the Method |  |  |
| General Policies |  |  |
| CLIA Certificate <insert a copy> |  |  |
| Method Validation |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| System Configuration |  |  |
| Policy |  |  |
| Procedure |  |  |
| Calibrating (Coding) |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Blood Sample Collection for testing on the ACCU-CHEK Inform II System |  |  |
| Performing Tests on the ACCU-CHEK Inform II System |  |  |
| Quality Control Testing |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Patient Testing |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Linearity Testing |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Proficiency Testing |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Results Review |  |  |
| Policy |  |  |
| Procedure |  |  |
| Transferring Data |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| ACCU-CHEK Inform II System Maintenance |  |  |
| Storing/Maintaining/Handling the ACCU-CHEK Inform II Meters and Base Units |  |  |
| Policy and Practice |  |  |
| Troubleshooting the ACCU-CHEK Inform II System |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Troubleshooting Operational Issues |  |  |
| Troubleshooting Questionable Results |  |  |
| Replacing ACCU-CHEK Inform II Meters |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| ACCU-CHEK Inform II System Elements of Quality Assurance |  |  |
| Policy |  |  |
| Process |  |  |
| Documentation |  |  |
| ACCU-CHEK Inform II System Operator Certification/Recertification |  |  |
| Policy |  |  |
| Process |  |  |
| Procedure |  |  |
| Operator In-Service - Policy, Process and Procedure |  |  |
| Operator In-Service Guide |  |  |
| ACCU-CHEK Inform II System In-Service Checklist |  |  |
| ACCU-CHEK Inform II System Test |  |  |
| ACCU-CHEK Inform II System Test Answer Key |  |  |
| Related Policies and Procedures |  |  |
| <Insert Blood Sample Collection Policies and Procedures> |  |  |
| <Insert Infection Control Policies and Procedures> |  |  |
| <Insert Applicable Patient and Employee Safety Policies and Procedures> |  |  |
| Appendix |  |  |
| <List and insert related and referenced documents here or indicate where each document can be found, such as: - ACCU-CHEK Inform II Test Strip Package Insert and MSD sheet - ACCU-CHEK Inform II Control Package Insert and MSD sheet - ACCU-CHEK Linearity Kit package insert and MSD sheet - ACCU-CHEK Inform II Operator’s Manual - Miscellaneous manufacturer bulletins and technical literature - Manufacturer evaluation protocols - Operator competency records> |  |  |

ACCU-CHEK Inform II Glucose Monitoring System

Comprehensive Policies, Processes and Procedures Manual Approval

- Policy:  It is the responsibility of the ____________________________ <enter the department and title of the staff position(s)> to compose and maintain the ACCU-CHEK Inform II System Policies, Processes and Procedures Manual.  This manual conforms to and is subject to the ________________________ <enter the name of your facility’s internal policy for documentation management and control.> It must be approved, signed and dated by _______________________ <Enter the title of the Director the program in your facility> when it is placed in use and then reviewed and approved annually throughout the use of the test system.  If a new Director assumes this responsibility, the procedures must be re-approved, signed and dated by the new Director. Changes to this manual, or discontinuation of the test system must be approved, signed and dated by the Director including the date of change or discontinuation. A copy of the manual with change history is maintained throughout the use of the test system and for at least ___ <enter the number of years you wish to retain this manual> years after it has been discontinued.

- Initial approval date: _____________ Date test discontinued: ______________

- Signature:_____________________________            Signature:_________________________

- Printed Name:__________________________           Printed Name:______________________

| - Date | - Reviewed/Revised | - By Whom (Signature) | - Title |
| --- | --- | --- | --- |

------------------------------------------------------------------------------------------------------------

|  | Extent of Use Policy |
| --- | --- |

The ACCU-CHEK Inform II system is used in the <enter the name of your facility/healthcare system> to quantitatively measure glucose in fresh venous, arterial and capillary (fingerstick and neonatal heelstick) whole blood. The system is used as an aid in monitoring the effectiveness of glucose control.

The system may be used on multiple patients when compliant with the cleaning and disinfecting recommendations of the FDA, CDC, and CMS found in the following documents:

- Use of Fingerstick Devices on More Than One Person Poses Risk for Transmitting Bloodborne Pathogens: Initial Communication

- CDC Clinical Reminder: Use of Fingerstick Devices on More than One Person Poses Risk for Transmitting Bloodborne Pathogens

--------------------------------------------------------------------------------------------

|  | Principle of the Method |
| --- | --- |

The ACCU-CHEK Inform II system quantitatively measures glucose in whole blood.  The enzyme on the test strip, mutant variant of quinoprotein glucose dehydrogenase from Acinetobacter calcoaceticus, recombinant in E. coli, converts the glucose in the blood sample to gluconolactone.  This reaction creates a harmless electrical DC current that the meter interprets for a glucose result.  The sample and environmental conditions are also evaluated using a small AC signal.

The system is calibrated with venous blood containing various glucose concentrations and is calibrated to deliver plasma-like results. The reference values are obtained using a validated test method. This test method is referenced to the hexokinase method and is traceable to an NIST standard.

------------------------------------------------------------------------------------------------------------

|  | Limitations of the Method |
| --- | --- |

- Policy: The ACCU-CHEK Inform II system is used within its limitations as described by the manufacturer’s most current literature on interference and limitations - See list below and manufacturer’s ACCU-CHEK Inform II test strip package insert and literature attached in the appendix section of this manual. <Insert current ACCU-CHEK Inform II test strip package insert and interference and limitations literature into the appendix section.>

The ACCU-CHEK Inform II test strips are for testing fresh capillary, venous, arterial, or neonatal whole blood. Cord blood samples cannot be used.

• Hematocrit should be between 10–65 %.

• Lipemic samples (triglycerides) in excess of 1800 mg/dL may produce elevated results.

• Blood concentrations of galactose >15 mg/dL will cause overestimation of blood glucose results.

• Intravenous administration of ascorbic acid which results in blood concentrations of ascorbic acid >3 mg/dL will cause overestimation of blood glucose results.

• If peripheral circulation is impaired, collection of capillary blood from the approved sample sites is not advised as the results might not be a true reflection of the physiological blood glucose level. This may apply in the following circumstances: severe dehydration as a result of diabetic ketoacidosis or due to hyperglycemic hyperosmolar non-ketotic syndrome, hypotension, shock, decompensated heart failure NYHA Class IV, or peripheral arterial occlusive disease.

• This system has been tested at altitudes up to 10,000 feet.

• The performance of this system has not been evaluated in the critically ill.

--------------------------------------------------------------------------------------------

|  | General Policies |
| --- | --- |

<Consider the list below and insert/delete/edit/add general policies according to your practices and preferences>

- The ACCU-CHEK Inform II blood glucose monitoring program resides under the authority of the _______________________________________ <enter the name of the department under which the CLIA certificate resides and attach a copy of the CLIA certificate at the end of this section.> The system is administered by the <enter the department and staff position that administers the program. Example: Laboratory Point-of-Care Coordinator.>   Policies and procedures related to the program are developed in cooperation and compliance with the policies of the following involved departments within the facility <consider the list below and insert/delete/edit/add all departments that are associated or consulted with at some level:

- Laboratory

- Nursing

- Administration

- Pharmacy

- Biomedical Engineering

- Materials Management

- Shipping/Receiving

- Information Technology

- Security/Safety

- Medical Records

- Add any other associated departments>

- The director of the program has determined that policies and procedures in this manual are in compliance with the ACCU-CHEK Inform II system manufacturer’s recommendations for proper product use as described in the most current versions of the ACCU-CHEK Inform II Operator’s Manual and all relevant reagent package inserts and related technical literature and bulletins.

- The most recent ACCU-CHEK Inform II Operator’s Manual provided by Roche Diagnostics is referenced for procedural information.  <Keep the current ACCU-CHEK Inform II Operator’s Manual in the Appendix section at the end of this manual or within easy access.>

- The most recent reagent package inserts provided by Roche Diagnostics are referenced for procedural information – See Appendix section of this manual. <Insert the most current ACCU-CHEK Inform II reagent package inserts in the appendix section.>

- ACCU-CHEK Inform II system supplies are stored in the ___ <enter the department where test strips, control solutions and replacement meters are stored.> Supplies on nursing units are replenished by <enter your policy for replenishing nursing unit supplies.>

- Reagents are not used past their expiration date.  ACCU-CHEK Inform II test strips expire on the date printed on the test strip vial label.  ACCU-CHEK Inform II control and linearity solutions expire on the date printed on the vial label, or 3 months from opening, whichever comes first.  Whenever an operator opens a vial of controls or linearity solution, he/she must handwrite the expiration date on the vial. That date will be either 3 months from opening or the date printed on the vial label, whichever comes first.

- Results Ranges: The normal, critical and reportable ranges of the ACCU-CHEK Inform II system are established as described in the table below.  Results that exceed the reportable and critical range of the system require follow up as described in the Patient Testing section of this manual.

- Reportable Range:  The meter is configured to display a non-numeric result               (RR LO/RR HI or LO/HI; measurement range default) for tests that exceed the established reportable range <Delete reference to “RRLO/RRHI” if you do not adjust the reportable range in the data management system.>

- Critical Range:  The meter is configured to display a non-numeric result (CR LO or CR HI) for tests that exceed the established critical range <Delete the previous sentence if you do not configure the meters to display non-numeric results when they exceed the critical range.>

| Hospital Unit (NICU, Adult, etc.) | Normal Range (mg/dL) | Critical Limits (mg/dL) | Reportable Range (mg/dL) |
| --- | --- | --- | --- |
|  | <Enter your selected normal range> | <Enter your selected critical limits> | <Enter your selected reportable range> |
|  | <Enter your selected normal range> | <Enter your selected critical limits> | <Enter your selected reportable range> |
|  | <Enter your selected normal range> | <Enter your selected critical limits> | <Enter your selected reportable range> |

- Comments:  <Enter your requirement for attaching comment(s) to test results, include a list of available pre-set comment(s) and a list of results that might initiate additional comment(s) from the operator. Example text: Appropriate comment(s) are required to be attached to the following result(s) and additional comments are optional by the operator(s) with the following results.>

- All results

- Control results that are out of range

- Patient results that exceed critical limits

- Patient results that exceed reportable limits

- Any test that an operator intends to repeat (questionable results, known procedural errors, etc.)

- Repeated test results

- <insert other occasions when you wish to require comment(s).>

- Operator Identification:  Operators performing tests are identified in the system by means of:  <Customize the items below>:

- ___ <Enter a description of the operator ID number format>

- Operator ID numbers <may/must> be entered by <enter your policy for requiring barcode scanning and/or manual entry.>

- Only a currently certified operator may perform testing on the ACCU-CHEK Inform II system. Operators must successfully complete the in-servicing program described in this manual.  Operators may perform tests only after entering their own ID and must under no circumstances provide their ID to enable another operator to perform testing. <enter any policy you may have regarding potential disciplinary action for testing under another operator’s ID.>

- Infection control and operator and patient safety: The following related policies and procedures are followed for purposes of infection control, patient and operator safety when conducting tests with the ACCU-CHEK Inform II system and are included in the Related Policies and Procedures section of this manual: <list all related policies and procedures.>

- Patient Sample Collection: Proper blood sample collection is an essential and integral part of bedside glucose testing.  All staff members who perform testing on the ACCU-CHEK Inform II system must be fully trained to collect samples according to established sample collection policies and procedures.  The following procedures are referenced in the patient testing section of this manual and are included in the Blood Sample Collection for testing on the ACCU-CHEK Inform II system section of this manual <List all of your established procedures for capillary (fingerstick and neonate heelstick), venous, arterial and line sample collection.>

- ------------------------------------------------------------------------------------------------------------

Method Validation

|  | Policy |
| --- | --- |

Method validation studies as outlined in the process table and procedure below are conducted prior to placing a new system into service for patient testing.  The purpose of the validation studies is to ensure that the ACCU-CHEK Inform II blood glucose monitoring system performs acceptably and meets the needs of the facility.  The ___<enter the title of the staff position(s) responsible for arranging and conducting the studies or collecting validation literature> is responsible for arranging, conducting and documenting the method validation studies and may enlist the council and assistance of manufacturer representatives. The results of the method validation studies must meet the expectations established in the Interpretation of Results section of the procedure prior to placing the system into service.

|  | Process Table |
| --- | --- |

Purpose:  This process describes the inter-related actions taken in order to validate the ACCU-CHEK Inform II System method.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| The performance of the ACCU-CHEK Inform II system is validated prior to placing the system into service. The system is validated by means of: <consider the list below and insert/delete/edit/add validation methods according to your practice and preference>: - Linearity on each or a subset of meter(s) - Meter-to-meter comparability with blood - Precision Studies - Accuracy Studies - Performance literature provided by manufacturer | <enter the staff position(s) responsible for arranging and executing method validation studies; e.g., Point-of-Care Coordinator, manufacturer’s technical representative> | ACCU-CHEK Inform II Method Validation - General Procedure (See below) <List all of the procedures utilized in conducting method validation studies. This may include procedures provided by the manufacturer.> |
| A report of the method validation studies is obtained/generated. | <enter the staff position(s) responsible for arranging and executing method validation studies: e.g., Point of Care Coordinator with assistance from manufacturer> | ACCU-CHEK Inform II Method Validation - General Procedure (See below) |
| Validation study results are evaluated for acceptability according to established criteria. <Insert actions taken if the results are acceptable or unacceptable> | <enter the title of the healthcare professional responsible for declaring the method validation study results acceptable or unacceptable> | ACCU-CHEK Inform II System Method Validation - General Procedure (See below) |
| The report of the method validation studies is stored in the Quality Assurance section of this manual. | <enter the staff position(s) responsible for storing the report> | ACCU-CHEK Inform II System Method Validation - General Procedure (See below) |

|  | General Procedure |
| --- | --- |

<Note: if you utilize system performance literature provided by the manufacturer as method validation you will not need a procedure section>

Purpose:  This procedure describes the general steps taken to validate the ACCU-CHEK Inform II blood glucose monitoring system.

Procedure:

- Ensure that all meters involved in method validation studies are in good working order: They must be: fully assembled, charged, coded to the correct test strip lot and documented QC to confirm operation

- Conduct the following tests/studies to validate the ACCU-CHEK Inform II system: <consider the list below and insert/delete/edit/add study types according to your established practices and preferences>

  - Control Level 1 (Low) and Control Level 2 (High)

  - Linearity testing on all or a subset of new meters

  - Precision studies

  - Accuracy Studies

  - Meter-to-Meter comparability studies

  - <Enter and describe any other types of method validation studies you wish to conduct.>

- Procedure for performing Quality Control testing: See the Performing Quality Control Testing procedure in this manual.

- Procedure for performing Linearity testing: See the Performing Linearity Testing procedure in this manual.

- Procedure for performing Precision Studies: <Either insert your own procedure for conducting precision studies or refer to the Roche Diagnostics Precision Studies Protocol for Glucose Meter Evaluations, place it in an appendix and reference the appendix number.>

- Procedure for performing Accuracy Studies: <Either insert your own procedure for conducting precision studies or refer to the Roche Diagnostics Accuracy Study Protocol for Glucose Meter Evaluations, place it in an appendix and reference the appendix number.>

- Procedure for performing Meter-to-Meter Comparability Studies <Either insert your own procedure for conducting meter-to-meter comparability studies or refer to the Roche Diagnostics Meter-to-Meter Comparability Study Protocol for Glucose Meter Evaluations, place it in an appendix and reference the appendix number.>

- Procedure for performing <Enter any other study types and the procedures here.>

- Generate a report of the data that produces the statistical analysis needed to assess acceptability as outlined in the Interpretation of Result section of this procedure.

- Submit the report to the __<enter the title of the staff position responsible for approving the results> for evaluation of the acceptability of the results.

Interpretation of Results:

- The following table contains the acceptability criteria for method validation study results:

| Study Type | Acceptability Criteria |
| --- | --- |
| Quality Control Testing | All QC results must be__________ <insert your preferred criteria.> |
| Linearity Control Testing | Linearity test results must be _____________<insert your preferred criteria> |
| Precision Studies | Within-Run/Inter-Meter/Inter-Day <select appropriate type> precision study results must show : Standard Deviation less than___________ <insert your preferred maximum SD> Coefficient of Variation less than________ <insert your preferred maximum CV> |
| Accuracy Studies | Accuracy Study results must ___________ <insert your preferred acceptability criteria> |
| Meter-to-Meter Comparability Study | Results must show : Standard Deviation less than___________ <insert your preferred maximum SD> Coefficient of Variation less than________ <insert your preferred maximum CV> |
| Other Studies <Name other study types> | <insert your preferred acceptability criteria for other studies> |

- If the outcome of the method validation studies is within acceptability criteria and appropriate signatures obtained, the system may be placed into service.

- If the system does not meet acceptance criteria, call to consult with ACCU-CHEK Customer Care to investigate questionable results at 1-800-440-3638.

- Place a copy of the method validation report in an appendix at the end of this manual.  Retain the report for the time of use of the test system and for <enter your preferred number of years> years following discontinuation of the test system.

System Configuration

|  | Policy |
| --- | --- |

ACCU-CHEK Inform II meters are delivered with preconfigured default settings.  Various parameters of the ACCU-CHEK Inform II system can be customized by configuration through the ___ <enter name of your data management system> software or through the set-up menu on the meter.  The system is configured through the <describe whether you configure your system through your data management system or the meter set-up functions> to reflect <enter the name of your healthcare facility/system> practices before it is initially placed into service and may be modified as needs or practices change.  Initial configuration of the ACCU-CHEK Inform II system and all subsequent modifications are the responsibility of ___<enter the title of the position responsible for initial system configuration.>

The <enter the name of your healthcare facility/system> has selected custom settings as outlined in the System Configuration Report attached at the end of this section: <generate a system configuration report from your data management system and insert it at the end of this section.  Update the report whenever you change the system configuration.>

|  | Procedure |
| --- | --- |

Refer to the ACCU-CHEK Inform II Operator’s Manual and <insert the name of the data management system manual> for system set-up / configuration instructions.

Calibrating (Coding)

|  | Policy |
| --- | --- |

Each box of test strips contains a code key.  Each code key belongs to a single lot and provides important information about the lot-specific properties of the ACCU-CHEK Inform II test strip.  The properties of each lot number of test strips are downloaded (as a code file) from the code key into the ACCU-CHEK Inform II system by means of the code key reader. A code file is uploaded into the ACCU-CHEK Inform II system for every test strip lot that is received by the <enter the name of your healthcare facility/system.>  <Insert your policy on how to upload code file information such as meter by meter or via centralized distribution via a data management system.> <Insert any policy you may have for changing the editable parameters of the code file, such as shortening the expiration date or changing the control ranges.>  The code file for each in-use test strip lot resides in all meters so that end users on nursing units are able to access and select the correct test strip lot for testing.  New test strip codes are uploaded into the system <enter your policy regarding what staff position(s) is/are permitted to add new reagent lots into the system.>

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to calibrate the ACCU-CHEK Inform II system to a new test strip lot.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| A new test strip lot arrives for use in the <enter the name of your healthcare facility/system.> | <Enter the department(s) staff position(s) for receiving and implementing new test strip lots.> | <Enter the title of any procedure that contains reference to your internal policies and procedures for receiving and implementing new product.> Calibrating (Coding) the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual |
| The new test strip lot is entered into a single ACCU-CHEK Inform II meter by means of the code key reader. | Point of Care Coordinator <List any other staff position(s) you intend to allow to enter test strip lots into the system.> | Calibrating (Coding) the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual |
| <Delete this row in the table if you do not use a data management system.> The meter that has received the code file for the new test strip lot is downloaded into the <enter the name of your data management system.> | Point of Care Coordinator | Calibrating (Coding) the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual <Enter the title of your data management system operator’s manual or user’s guide> |
| <Delete this row in the table if you do not use a data management system.> The code file is disseminated to all remaining meters by means of connectivity with the <enter the name of your data management system> | Point of Care Coordinator | Calibrating (Coding) the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual <Enter the title of your data management system operator’s manual or user’s guide> |

|  | Procedure |
| --- | --- |

Purpose:  The purpose of this procedure is to give basic instructions for uploading a new test strip code file into the ACCU-CHEK Inform II system meters using the code key reader.  The instructions in the procedure are abbreviated.  Refer to the ACCU-CHEK Inform II Operator’s Manual for detailed and pictorial instructions and follow your data management instructions to disseminate code key information to all meters, if applicable.

Materials: Gather the following materials in preparation for uploading a new test strip lot code file into the ACCU-CHEK Inform II system:

- A single ACCU-CHEK Inform II meter – Fully charged

- A code key from the new strip lot

- A code key reader

- A downloading (network connected) base unit

- <Enter the name of your data management system>

Procedure:

- Turn on the ACCU-CHEK Inform II meter.

- Enter or barcode scan your operator ID <insert your policy for entering operator ID’s; e.g., barcode scanning and/or manual entry> NOTE: If the operator ID you enter is not accepted, attempt to re-enter it.  If it is still rejected, contact your supervisor or Point of Care Coordinator.  DO NOT attempt to perform tests under another operator’s ID.

- From the Main Menu, touch the forward arrow button to open the Main Menu 2 screen.

- Touch Strip Lots to open the test strip lot code file menu.

- Touch Add if you want to add the information for a new test strip lot from a new code key. The Add Strip Lot screen opens.

- Insert the new code key in the opening of the code key reader. A LED starts flashing green to signal that the code key reader is ready to transfer data.

- Place the code key reader on a level surface such as a bench or table

- Hold the meter 4-6 in (10-15 cm) above the code key reader so that a connection can be made between the infrared window on the bottom of the meter and the infrared window on the top of the code key reader.

- Touch the forward arrow key to begin downloading data.

- Successful transmission will display two progress messages:  “Please Wait – Connecting to code key reader and “Please Wait – Receiving Code Key Contents.”

- A Strip Lot Confirmation screen will ask you to confirm use of the suggested values for the test strip lot

- Touch the  button to store the data for this test strip lot number in the meter without changes, or Touch the  button to modify the data for this test strip lot number (see the ACCU-CHEK Inform II Operator’s manual for detailed instructions on editing parameters of the test strip lot.)

- You will then see a Make ‘Current’ screen asking you if you want to make the test strip lot that you are entering the current test strip lot.

- Touch the  button to confirm that you want this lot number to be the lot number currently in use, or touch the  button to store the entries without making the lot number the current lot number.

- Touch the Main Menu icon in the center of the bottom of the screen to return to the Main Menu.

- Dock the meter in the network connected base unit to send the new test strip lot information from the meter into the data management system for centralized test strip lot distribution.

Blood Sample Collection

|  | Policy |
| --- | --- |

- Proper blood sample collection is an essential and integral part of bedside blood glucose testing.  All staff members who perform testing on the ACCU-CHEK Inform II system must be fully trained to collect samples according to established sample collection policies and procedures.  The following procedures are referenced in the patient testing section of this manual and are included in this section <List all of your established procedures for capillary (fingerstick and neonate heelstick), venous, arterial, auto-disabling single-use lancing device and line sample collection.>

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to collect high quality blood samples for testing on the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| Policies and procedures for collecting blood samples for testing on the ACCU-CHEK Inform II meters are established. | <Enter the titles of the staff positions responsible for establishing blood sample collection policies and procedures.> | <Enter the titles of all relevant blood collection policies established in your facility.> |
| Staff members responsible for ACCU-CHEK Inform II testing are trained to collect samples according to established policies. | <Enter the titles of staff positions responsible for training staff in blood sample collection procedures.> |  |
| Properly trained staff members collect blood samples for testing according to established policies and procedures | All operators who are currently certified to perform meter testing. |  |

|  | Patient Samples and Sample Collection |
| --- | --- |

Acceptable Samples:

- The following fresh whole blood sample types may be used:

  - Venous whole blood

  - Arterial whole blood

  - Capillary (non-neonate fingerstick and neonate heelstick) whole blood

- The following anticoagulants are acceptable (do not use any other anticoagulants for meter testing):

  - Lithium or Sodium Heparin; EDTA

|  | Blood Sample Collection Procedures |
| --- | --- |

<Insert your facility’s procedures for collecting venous, arterial, line sample, capillary fingerstick and heelstick blood samples for ACCU-CHEK Inform II testing here.  The following procedure is offered as an example for fingerstick capillary blood sample collection.  Example procedures for infant heelstick, venous, arterial  and line draw sample collection are not offered because such procedures are specific to the equipment used in each facility and generally require advanced qualification and training.>

Fingerstick Capillary Blood Sample Collection Procedure (Example)

NOTE:  Follow all facility safety and infection control policies when collecting blood samples when collecting blood samples.

Fingerstick Sample Collection

- Assemble the materials you will need to collect a blood sample (gloves, skin preparation pad, auto-disabling single-use lancet device, gauze or cotton ball).

- Wash hands and don gloves and any other personal protective equipment as required by infection control and isolation policies and procedures.

- Follow all facility infection control protocols when testing isolation patients <insert appropriate isolation procedure information here according to your facility’s policies.>

- Assess the patient for compromised peripheral blood flow.  Fingertips should be warm and pinkish when the hand is gently massaged from the palm outward to the fingertips.  Fingertips should not appear pale, bluish or mottled.  Patients with compromised peripheral blood flow are not good candidates for fingerstick testing.

- Select the finger site for puncture.  It is preferred to select the side of a middle or ring finger that has not been punctured recently.

- Enhance blood flow to the selected puncture site by means of:

  - Warming the intended puncture site

  - Instructing the patient to flex and move the arm, wrist, hand and fingers while you are assembling your supplies and preparing the system for testing

  - Positioning the intended puncture site below heart level

  - Gently massaging in an outward (distal) direction from the palm and the base of the finger to the fingertip.

- Cleanse the puncture site by means of <insert your facility’s policy for cleansing the puncture site.>  Allow the site to air dry completely before puncturing.

- Advise the patient of imminent puncture.

- <Insert device specific instructions for puncturing the site with the device in use in your facility. Example text using ACCU-CHEK® Safe-T-Pro Plus lancet:

  - Twist off the protective cap of the Safe-T-Pro Plus lancet and discard

  - Choose the desired depth setting

  - Hold the Safe-T-Pro Plus lancet tip against the puncture site

  - Press the purple trigger button

  - Withdraw the Safe-T-Pro Plus lancet from the site>

- Hold the puncture site downward and gently apply intermittent proximal to distal pressure along the finger toward the puncture site to express a blood drop.  Do not apply strong repetitive pressure at the fingertip as it may cause hemolysis or contamination of the sample with tissue fluid and may lead to questionable results.

- <Insert your policy regarding wiping the first drop away. This is advantageous because it ensures that the cleansing agent is dry, it stimulates blood flow and clears interstitial fluid from the sample,8>

- Apply a well-formed drop to the ACCU-CHEK Inform II test strip as described in the patient testing procedure that follows.

- Apply gentle direct pressure to the puncture site for several minutes and elevate the hand to reduce blood flow to the fingertip.  Check the site to ensure that it is no longer bleeding before leaving the patient bedside.

- Discard all sample collection and testing materials by means of <insert your facility’s policy for safely discarding sharps and testing materials.>

- Wash hands before leaving the patient room.

- Clean and disinfect according to facility policy.  See System Maintenance section below for detailed instructions on cleaning and disinfecting.

NOTE: The FDA recommends that POC testing devices should be used only on one patient and not shared.  If dedicating blood glucose meters to a single patient is not possible, the meters must be properly cleaned and disinfected after every use

Performing Tests on the ACCU-CHEK Inform II System:

Quality Control Testing

|  | Policy |
| --- | --- |

Quality control testing is performed as a primary means of ensuring on-going proper performance of the ACCU-CHEK Inform II system.  The process table and procedure below describe the steps taken to perform quality control testing on the ACCU-CHEK Inform II system.  Low and high quality control testing is performed on the following occasions:

- ____ time(s) per 24 hour period of patient testing <Enter your preferred frequency of control testing.>

- When a new test strip vial is opened and placed into service.

- <List other times you expect quality control testing to be performed – e.g.: whenever you observe unexpected results of any kind; when the test strip lot is changed; when a meter is dropped, etc.>

- <Enter your policy for allowing STAT test by-passing of scheduled quality control testing.>

- Acceptable quality control results are defined as results that are within <enter your preferred acceptable limits; e.g.: within manufacturer’s posted limits or whatever limits you choose to establish internally.  Note: If you choose to establish internal limits, include a description of how those limits are derived.  It is not advised to set limits that exceed the manufacturer’s posted limits.>

- Entering new control lots into the ACCU-CHEK Inform II system <Insert your policy regarding who is authorized to enter new control lots.  Refer to the ACCU-CHEK Inform II Operator’s Manual or data management system user’s manual for instructions on entering new control lots.>

Quality control testing results must be acceptable prior to performing patient testing, according to the Interpretation of Results section of the procedure below.

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to conduct Quality Control testing on the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| A policy and procedure for quality control testing is established. | <Enter the title of the health care professional responsible for the hospital bedside blood glucose monitoring system> | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert ACCU-CHEK Inform II control solutions package insert <Enter the title of any other internal procedure that contains reference to frequency of quality control testing in the laboratory> |
| Qualified operators perform quality control testing | All operators who are currently certified to perform meter testing (Note: quality control testing is part of the operator certification process; therefore, an operator may perform quality control testing prior to certification in this context, but may not perform quality control testing as part of the routine usage of the system until he or she is formally certified.) | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert ACCU-CHEK Inform II control solutions package insert <Enter the title of any other internal procedure that contains reference to the procedure for ACCU-CHEK Inform II quality control testing> |
| Quality control results are interpreted. | Any operator who is currently certified to perform quality control testing or his/her designated supervisor. | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure ACCU-CHEK Inform II System Configuration Policy, Process and Procedure Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert <Enter the title of any other internal procedure or document that contains reference to the policies regarding interpretation of quality control testing results.> |
| Quality control results are downloaded into the electronic data management system. | Any operator who is currently certified to perform quality control testing or his/her designated supervisor. | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Transferring Data from the ACCU-CHEK Inform II System – Purpose, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert <Enter the title of any other internal procedure that contains reference to the policies regarding recording of quality control testing results.> |

|  | Procedure |
| --- | --- |

Purpose:  This procedure describes the basic steps taken to perform quality control testing on the ACCU-CHEK Inform II system.  Refer to the ACCU-CHEK Inform II Operator’s Manual for more detailed pictorial instructions.

Samples:  ACCU-CHEK Inform II low and high glucose controls are used for quality control testing

Materials:  Gather the following supplies in preparation for quality control testing:
Note: <Insert a statement regarding where supplies are stored and how they are obtained to replenish stock on nursing units;  Example text: “Contact the Laboratory if you need to replenish your stock of ACCU-CHEK Inform II test strips and controls solutions or to replace a non-operational meter.>

- ACCU-CHEK Inform II meter – fully charged and coded to the test strip lot you intend to use.

- ACCU-CHEK Inform II Control Level 1 (Low) and Control Level 2 (High), the lot and range information will be available in meter.

- ACCU-CHEK Inform II test strip vial

- NOTE about reagents:  Reagents are not to be used past their expiration date.  ACCU-CHEK Inform II strips expire on the date printed on the strip vial label. ACCU-CHEK Inform II Control and linearity solutions expire on the date printed on the vial label, or 3 months from opening, whichever comes first. Whenever an operator opens a vial of controls or linearity solution, he/she must handwrite the expiration date and his/her on the vial. That date will be either 3 months from opening or the date printed on the vial label, whichever comes first.

Procedure:

- Turn on the ACCU-CHEK Inform II meter.

- Enter or your operator ID by means of <insert your policy for entering operator IDs; e.g., barcode scanning and/or manual entry> NOTE: If the operator ID you enter is not accepted, attempt to re-enter it.  If it is still rejected, contact your supervisor or Point of Care Coordinator.  DO NOT attempt to perform tests under another operator’s ID.

- From the Main Menu, touch Control Test.

- Select the level of control that you wish to test <insert instructions on whether the control level is to be entered by barcode scanning and/or manual entry>

- Confirm that the meter is coded (calibrated) to the same test strip code that is printed on the test strip vial by <enter your policy for confirming the test strip code; e.g., barcode scanning and/or screen confirmation>. Contact your supervisor or Point-of-Care Coordinator if you are unable to confirm the correct test strip code.

- The meter will display a picture of a test strip with a downward flashing arrow on the meter indicating that you are ready to insert a test strip into the meter. Remove a test strip from the vial and immediately recap the vial. Insert the test strip into the meter in the direction of the arrows and with the “ACCU-CHEK” lettering facing upward. The meter will display a flashing drop above the test strip icon when the test strip is properly inserted indicating that you are ready to apply control solution.

- Apply control solution to the front edge of the test strip. The solution will fill the yellow sample chamber by capillary action. Do not apply sample to the top of the test strip. Once sufficient sample has been detected, the measurement begins. An hourglass icon indicates that the measurement is in progress. You will get an error message if the sample is insufficient. If this occurs, you will need to repeat the test.

- The measurement is complete when the result is displayed on the meter screen.  See the “Interpretation of Results” and “Troubleshooting” sections below for guidance on what to do if your result is “Fail” or shows an “out of range” message associated with the result.

- Remove the test strip and dispose of it ___________<insert your desired method of disposal.>

- Touch the comment button ( ) to enter an appropriate comment(s) as required in the “General Policies” section of this manual.

- Touch the   button to confirm the result and send the result from the meter wirelessly or place the meter in the base unit to send the result and record the result into the electronic data management system. The base unit also charges the meter.

|  | Interpretation of Quality Control Results |
| --- | --- |

- Results are displayed on the screen as ____ <enter either “a numeric value” or “Pass or Fail” depending upon how your system is configured.> Any result that shows an “out of range” message or “Fail” is an indication that the system may not be performing correctly for patient testing.

- Patient testing may not be performed if quality control testing results are not within acceptable limits and the meter will not display the patient testing option if scheduled quality control results exceed acceptable limits.

|  | Troubleshooting |
| --- | --- |

Call ACCU-CHEK Customer Care at 1-800-440-3638 for assistance any time you have questions or concerns regarding the ACCU-CHEK Inform II system.

- Troubleshooting meter operational issues: <enter a description of the extent of troubleshooting you intend to allow to take place on the nursing unit. Example text: “If the meter fails to function at any point in the procedure or if you get an error message associated with the result, make a note of the malfunction or error message and attempt to repeat the test.  If the error persists, sequester the meter and test strip vial involved and deliver them to <enter the department and staff position(s) responsible for advanced troubleshooting.>

- If the error message “Strip Defect Error” appears on the display, the test strip may be defective or the blood glucose result may be extremely low and below the meter’s measurement range. Refer to the test strip package insert, perform a quality control test using a new test strip, review proper testing procedure, and repeat the blood glucose test, or follow your facility’s testing policy.

- If the meter displays “Type Bad Dose,” there may be insufficient amount of blood on the test strip. Repeat the test using a new test strip, ensuring proper sample application, or refer to the test strip package insert.

- Guidance for interpreting on-screen message and error codes:

| All error messages displayed by the system have a letter identifying the message type, a number and a description of the error to help the operator take action to resolve the problem. The different message types are in the table below. |  |
| --- | --- |
| - E | Identifies the notification as an Error. The information notifies the operator that an error has occurred. |
| - W | Identifies the notification as a Warning. The information does not block the operator from continuing, but rather gives the operator information that may suggest an alternate workflow is required. |
| - I | Identifies the notification as Informational only. Informational notifications present the operator with contextual information, and allow the operator to proceed after confirming the notification. |
| - D | Identifies a Decision point. Decision notifications provide the operator with a choice based on contextual information. |

- Troubleshooting out-of-range quality control results: Quality control results must be  _________<Enter your established limits.>  Take the following actions if your results are not within established limits:

<Enter your policy for following up on an out-of-range quality control

result.  Include information on where to take the meter and reagents if the operator is unable to resolve the issue.>

Example troubleshooting text:

- Add a comment(s) to the out-of-range result indicating that the test will be repeated.

- Repeat the test one time using the same test strip vial, control solution(s) and meter.

- If the result is in range, you may proceed to patient testing.

- If the repeat test using the same test strip vial is still out of range, repeat the test using a different test strip vial. If the result is within range, you may proceed with patient testing using the new test strip vial.  Discard the vial that failed quality control testing.

- If the repeat test using the different test strip vial remains out of range, repeat the test using a new vial of control solution. If the result is within range, you may proceed with patient testing. Discard the control vial that failed quality control testing.

- If the repeat test using the new control vial remains out of range, remove the meter and all of the test strips and controls you used from service and deliver them to the _________________department for further investigation <insert the department responsible for advanced troubleshooting.>

Patient Testing

|  | Policy |
| --- | --- |

Patient testing with the ACCU-CHEK Inform II system quantitatively measures glucose in fresh venous, arterial, neonatal heelstick and capillary whole blood from the finger and is used as an aid in monitoring the effectiveness of glucose control. The procedure below describes the steps taken to perform patient testing on the ACCU-CHEK Inform II system.  General patient testing policies are as follows:

- Patient testing must be ordered according to established test ordering policies and procedures.

- Patient samples are collected according to established phlebotomy procedures.  See the “Blood Sample Collection for testing on the ACCU-CHEK Inform II System” section of this manual.

- Patients are identified for testing by means of <enter the patient identification number and any other identifier you choose to use.> Patient identifiers are entered into the ACCU-CHEK Inform II system when prompted by the meter.

- Patient identification numbers are entered into the ACCU-CHEK Inform II system by means of <barcode scanning and/or manual entry and or both.>

- Special Patient Identification Conditions:<Insert your policy for entering patient identification under special circumstances, such as, ER, newborn infants or STAT tests on patients without an established ID; when performing practice testing, when patient ID is not accepted by the system.>

- Critical limits for patient testing are _________<enter your lower and upper critical limits and a description of whether they will be displayed numerically or as CRHI/CRLO>. Patient test results that exceed critical limits must be followed-up by <enter your policy for follow-up of patient results that exceed critical and/or reportable limits.>

- Reportable limits for patient testing are____ <enter your lower and upper reportable limits.> Results that exceed the reportable limits of the system will be displayed as RR HI/RR LO or HI/LO, the measurement range default <delete reference to RR HI/RR LO if you do not customize the reportable range in the data management system.> Patient test results that exceed reportable limits must be followed-up by <enter your policy for follow-up of patient results that exceed critical and/or reportable limits.>

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to conduct patient testing on the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| A patient bedside glucose test is ordered according to established policy for patient testing Exception: <Describe any policy you have established that allows for unordered STAT testing to be performed as a critical intervention in clinical emergencies> | <Enter the title of the healthcare professionals responsible for ordering tests: example – Physician; nursing> | <Enter the name of the procedure established for proper test ordering in your facility> |
| A properly trained staff person collects a blood sample appropriate for testing on the ACCU-CHEK Inform II System. | <Enter the staff title(s) qualified and responsible for collecting capillary (fingerstick and neonate heelstick) arterial, venous and line draw blood samples> | - <List all of your established phlebotomy procedures for capillary (fingerstick and neonate heelstick), venous, arterial and line sample collection.> - <List all internal policies and procedures related to infection control and operator and patient safety> |
| A properly certified operator performs the test | Any operator who is currently certified to perform testing | Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert |
| Patient test results are interpreted for acceptability and/or need for follow-up according to established policy. | <Enter the staff title(s) qualified to interpret results; example: Any RN who is currently certified to perform patient testing or his/her designee.> | Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure ACCU_CHEK Inform II System Configuration Policy, Process and Procedure Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II General Policies ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert <Enter the title of any other internal procedure or document that contains reference to the policies regarding interpretation of patient testing results> |
| <Delete this row in the table if this does not apply to your facility.>Proper follow-up action is taken depending upon the interpretation of results | <Enter the staff title(s) qualified to interpret results; example: Any RN who is currently certified to perform patient testing or his/her designee.> | Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert <Enter the title of any other internal procedure or document that contains reference to the policies regarding follow-up of patient testing results.> |
| Patient test results are downloaded into the electronic data management system. | Any operator who is currently certified to perform testing or his/her designated supervisor. | Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure Transferring Data from the ACCU-CHEK Inform II System – Purpose, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert <Enter the title of any other internal procedure that contains reference to the policies regarding recording of patient testing results.> |

|  | Procedure |
| --- | --- |

Purpose:  This procedure describes the basic steps taken to perform patient testing on the ACCU-CHEK Inform II system and as instructed clean and disinfect the meter between each patient test. Refer to the ACCU-CHEK Inform II Operator’s manual for more detailed instructions. Tips for optimizing conditions for reliable results are included in the Procedure Notes section below.

NOTE:  Follow all facility safety and infection control policies when collecting blood samples. This procedure does not give detailed instructions for infection control and other patient and operator safety practices, which fall, under the scope of established policies and procedures for infection control and safety. See the Related Policies and Procedures section of this manual for detailed infection control and safety instructions.

Performing an ACCU-CHEK Inform II Patient Test

Materials: Gather the following supplies in preparation for patient testing:

Note: <Insert a statement regarding where supplies are stored and how they are obtained to replenish stock on nursing units; Example text: “Contact the Laboratory if you need to replenish your stock of ACCU-CHEK Inform II test strips and controls solutions or to replace a non-operational meter.”>

- ACCU-CHEK Inform II meter – fully charged and coded to the test strip lot you intend to use.

- ACCU-CHEK Inform II test strip vial.  Note about reagents:  Reagents are not to be used past their expiration date. ACCU-CHEK Inform II strips expire on the date printed on the strip vial label.

- Supplies for collecting a blood sample as required by established phlebotomy procedures, such as cotton balls, alcohol preps, approved auto-disabling single-use lancing device, (See appropriate sample collection procedures).

- Gloves

- Biohazard and sharps container

- Transfer pipette or syringe as needed if testing a venous, arterial or line draw sample.

Acceptable Samples:

- The following fresh whole blood sample types may be used:

- Venous whole blood

- Arterial whole blood

- Capillary (non-neonate fingerstick and neonate heelstick) whole blood

- The following anticoagulants are acceptable (do not use any other anticoagulants for meter testing):

- Lithium or Sodium Heparin

- EDTA

Procedure:

- Carefully assess the patient for any indication that bedside glucose testing may not be appropriate (See Procedure notes for patient assessment guidance). Consider the following potential interferences and compromising conditions:

• Hematocrit should be between 10–65 %.

• Lipemic samples (triglycerides) in excess of 1800 mg/dL may produce elevated results.

• Blood concentrations of galactose >15 mg/dL will cause overestimation of blood glucose results.

• Intravenous administration of ascorbic acid which results in blood concentrations of ascorbic acid >3 mg/dL will cause overestimation of blood glucose results.

• If peripheral circulation is impaired, collection of capillary blood from the approved sample sites is not advised as the results might not be a true reflection of the physiological blood glucose level. This may apply in the following circumstances: severe dehydration as a result of diabetic ketoacidosis or due to hyperglycemic hyperosmolar non-ketotic syndrome, hypotension, shock, decompensated heart failure NYHA Class IV, or peripheral arterial occlusive disease.

• This system has been tested at altitudes up to 10,000 feet.

• The performance of this system has not been evaluated in the critically ill.

- Take the meter and testing supplies to the patient location.

- Wash hands and don personal protective equipment (gloves, gowns, etc.) as required by infection control and isolation policies and procedures.

- Greet and identify the patient according to <insert your facility’s policy for proper patient identification.>

- Explain the procedure to the patient.

- Turn on the ACCU-CHEK Inform II meter.

- Enter  your operator ID by means of <insert your policy for entering operator ID’s; e.g., barcode scanning and/or manual entry> NOTE: If the operator ID you enter is not accepted, attempt to re-enter it. If it is still rejected, contact your supervisor or Point of Care Coordinator.  DO NOT attempt to perform tests under another operator’s ID.

- From the Main Menu, touch Patient Test.

- Enter the patient identification in the ACCU-CHEK Inform II system <insert your policy for entering patient IDs; e.g., barcode scanning and/or manual entry.>

- Confirm that the meter is coded (calibrated) to the same test strip code that is printed on the test strip vial by <enter your policy for confirming the test strip code; e.g.: barcode scanning and/or screen confirmation.> Contact your supervisor or Point-of-Care Coordinator if you are unable to confirm the correct test strip code.

- You will now see a picture of a test strip with a downward flashing arrow on the screen indicating that you are ready to insert a test strip into the meter.

- Remove a test strip from the vial and immediately recap the vial. Insert the test strip into the meter in the direction of the arrows and with the “ACCU-CHEK” lettering facing upward.  The meter will display a flashing drop above the test strip icon when the test strip is properly inserted indicating that you are ready to apply a blood sample.

- Collect an acceptable blood sample according to your facility’s established procedures (See the “Procedure Notes” section below for tips on optimizing sample quality)

- Fingerstick or neonate heelstick samples: Test immediately as the sample is collected.

- Venous, arterial or line draw samples: Test as soon as possible and no later than 30 minutes following collection. Be sure they are well mixed and that line draw samples have been thoroughly cleared of line fluids. Do not allow bubbles to enter the test strip-sampling chamber.

- <Insert your policy regarding wiping the first drop away when testing capillary samples – this is advantageous because it ensures that the cleansing agent is dry, it stimulates blood flow and clears interstitial fluid from the sample,,.>

- Apply blood to the front edge of the test strip. The sample will fill the yellow sample chamber by capillary action. Do not apply sample to the top of the test strip.

- Once sufficient sample has been detected, the measurement begins. An hourglass icon indicates that the measurement is in progress.

- After the sample has been obtained, apply gentle pressure to the puncture with a clean gauze square or cotton ball site for several minutes. If the patient is conscious and capable, enlist the patient’s assistance with applying pressure.

- The measurement is complete when the result is displayed on the screen. Depending upon how high or low the result is, it may appear in a numeric or non-numeric format.  See Interpretation of Results section below for interpretation of each result format.

- Remove the test strip and dispose of it ___<insert your desired method of disposal.>

- Touch  to enter up to three appropriate comment(s) as required in the “General Policies” section of this manual.

- Touch the  button to confirm the result and send the result from the meter wirelessly or place the meter in the base unit to send the result and record the result into the electronic data management system. The base unit also charges the meter.

- Document the result in the patient chart by means of___<insert your policy for documentation of patient results.>

- Follow up on any results that exceed critical or reportable limits according to policy.

- Clean and disinfect as necessary following your hospital’s policy

|  | Interpretation of Patient Results |
| --- | --- |

- Results may appear in any of the following formats and may require follow-up:

  - A numeric value

  - “HI” or “LO” meaning that the result is above or below the upper or lower reading limits of the ACCU-CHEK Inform II system (>600 mg/dL  or <10 mg/dL) <insert your policy for follow up of “HI” or “LO” results.>

  - “RR HI” or “RR LO” meaning that the result is above or below your facility’s established reportable limits which are: <insert your established reportable limits.  Delete this section if you do not customize the reportable range in the data management system.>

  - “CR HI” or “CR LO” meaning that the result exceeds your established critical limits which are:  <insert your established reportable limits.>

2.	Follow up on results that exceed critical or reportable limits as outlined in the policy section above.

|  | Troubleshooting |
| --- | --- |

Call ACCU-CHEK Customer Care at 1-800-440-3638 for assistance any time you have questions or concerns regarding the ACCU-CHEK Inform II system.

- Troubleshooting meter operational issues: <Enter a description of the extent of troubleshooting you intend to allow to take place on the nursing unit.  Example text: “If the meter fails to function at any point in the procedure or if you get an error message associated with the result, make a note of the malfunction or error message and attempt to repeat the test.  If the error persists, sequester the meter and test strip vial involved and deliver them to <enter the department and staff position(s) responsible for advanced troubleshooting.”>

- If the error message “Strip Defect Error” appears on the display, the test strip may be defective or the blood glucose result may be extremely low and below the meter’s measurement range.  Refer to the test strip package insert, perform a quality control test using a new test strip, review proper testing procedure, and repeat the blood glucose test, or follow your facility’s testing policy.

- If the meter displays “Type Bad Dose,” there may be insufficient amount of blood on the test strip. Repeat the test using a new test strip, ensuring proper sample application, or refer to the test strip package insert.

- Guidance for interpreting on-screen message and error codes:

| All error messages displayed by the system have a letter identifying the message type, a number and a description of the error to help the operator take action to resolve the problem. The different message types are in the table below. |  |
| --- | --- |
| - E | Identifies the notification as an Error. The information notifies the operator that an error has occurred. |
| - W | Identifies the notification as a Warning. The information does not block the operator from continuing, but rather gives the operator information that may suggest an alternate workflow is required. |
| - I | Identifies the notification as Informational only. Informational notifications present the operator with contextual information, and allow the operator to proceed after confirming the notification. |
| - D | Identifies a Decision point. Decision notifications provide the operator with a choice based on contextual information. |

- Troubleshooting questionable results: <Enter a description of the extent of troubleshooting you intend to allow to take place on the nursing unit.  Example text:

    - Consider whether the result is consistent with the patient’s history and clinical presentation.  Take the following action if you question the reliability of the result for any reason:

      - Add a comment(s) to the result indicating that the result is in question.

      - Perform quality control testing using the same meter and test strips.

      - If quality control test results are within range, repeat the patient test using the same test strips and meter.

      - If the quality control tests are not within range, sequester the meter and test strip vial involved and deliver them to the Point of Care Coordinator for advanced troubleshooting.

      - Repeat patient testing using another meter and test strip vial that have passed routine quality control testing.

|  | Patient Testing Procedure Notes |
| --- | --- |

- Good sample quality is essential for reliable results.  The following tips will help you to optimize the quality of a fingerstick capillary sample:

  - Assess the patient for signs of reduced peripheral circulation such as cold hands, pale, mottled or bluish nailbeds and administration of vasopressor medications.  If peripheral circulation is impaired, collection of capillary blood from approved sites is not advised, as the results might not be a true reflection of the physiologic blood glucose level.  This may apply in the following circumstances: severe dehydration as a result of diabetic ketoacidosis, or due to hyperglycemic hyperosmolar non-ketotic syndrome, hypotension, shock, decompensated heart failure NYHA class IV, or peripheral arterial occlusive disease.

  - Always take basic steps to stimulate blood flow to the intended puncture site.  Even if the patient’s peripheral circulation is not impaired this will help you to procure a sufficient and free flowing sample, especially in patients who are sedentary.  Common methods for stimulating blood flow include:

    - Warming the site

    - Positioning the hand below the heart

    - Gently massaging the hand by stroking from the palm outward to the fingertip

    - Asking the patient to move and flex his/her arm, wrist, hand and fingers while you are gathering supplies and preparing the system for testing

  - Be sure to allow the cleansing agent to completely dry before puncturing the intended site.

  - <Insert your policy regarding wiping the first drop away – this is advantageous because it ensures that the alcohol is dry, it stimulates blood flow and clears interstitial fluid from the sample,>

  - Do not attempt to express blood from a previous puncture site even if it still appears to be bleeding because it is likely that the clotting process has begun and may alter the sample matrix.  A fresh puncture should be made for every test.

- Do not place the meter on bedding after applying blood to the test strip.  This can cause the blood to be wicked out of the test strip by the bedding.

- Recap the test strip vial immediately after you remove the test strip you plan to use.  Never use test strips found loose outside the vial.  Discard any vial of strips that you find uncapped.

Linearity Testing

|  | Policy |
| --- | --- |

Linearity testing, also known as calibration verification, is a means of testing the ACCU-CHEK Inform II system’s ability to give correct readings through a broad range of glucose levels.  It is used as a periodic check on system performance.  Linearity testing is conducted according to the following schedule: <enter your preferred schedule for linearity testing – include occasions, such as when a system is placed into service, as part of troubleshooting incidents of questionable meter performance, every six months, when a new test strip lot is placed into service etc.  Also include how many meters, how many levels and how many replicates of each level you prefer to test.>

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to conduct linearity testing on the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| A policy and schedule for performing linearity testing is established | <Enter staff position(s) responsible for developing the policies and schedule related to linearity testing> | Performing Tests on the ACCU-CHEK Inform II System – Linearity Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual <Enter the title of any other internal procedures that contain reference to the policies regarding linearity testing> |
| A properly trained Point-of-Care Coordinator or his/her designee performs linearity testing and downloads the results into the data management system. | <Insert the staff position(s) responsible for administering the program> | Performing Tests on the ACCU-CHEK Inform II System – Linearity Testing Policy, Process and Procedure Transferring Data from the ACCU-CHEK Inform II System – Purpose, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert ACCU-CHEK Linearity Kit package insert |
| A report of the linearity testing is generated by the data management system and the report is retained according to established record keeping policies. | <Insert the staff position(s) responsible for administering the program> | Performing Tests on the ACCU-CHEK Inform II System – Linearity Testing Policy, Process and Procedure <Enter the title of the data management system operator’s manual> ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Linearity Kit package insert |

|  | Procedure |
| --- | --- |

Purpose:  The purpose of this procedure is to provide basic instructions for the system administration staff to perform linearity testing.  The steps in this procedure are abbreviated.  Refer to the ACCU-CHEK Inform II Operator’s Manual for more detailed instruction.

Acceptable Samples: The ACCU-CHEK Linearity Kit is used to perform linearity testing

Entering new linearity kit lots into the ACCU-CHEK Inform II System: <insert your policy regarding who is authorized to enter new linearity kit lots.  Refer to the ACCU-CHEK Inform II Operator’s Manual or data management system user’s manual for instructions on entering new linearity kit lots.>

Materials:  Gather the following supplies to perform linearity testing:

- ACCU-CHEK Inform II meters – fully charged and coded to the test strip lot you intend to use

- ACCU-CHEK Inform II test strips

- ACCU-CHEK Linearity Kit – loaded for use in the meters to be tested.

Procedure:  See the ACCU-CHEK Inform II Operator’s Manual for complete instructions on performing linearity testing.

Interpretation of Results:

- See the ACCU-CHEK Inform II Operator’s Manual for guidance on interpreting linearity testing <or insert your preferred policy for interpreting linearity testing results.>

- Results are acceptable if they are <enter your preferred limits of acceptability; e.g.: within manufacturer limits.>  Call ACCU-CHEK Customer Care at 1-800-440-3638 if your results are outside allowable limits.

Procedure Notes :

- <Identify where you store linearity-testing records.  Suggestion:  Consider storing your linearity records in the Quality Assurance section of this manual.>

Proficiency Testing

|  | Policy |
| --- | --- |

Participation in proficiency testing is part of the quality assurance program to ensure proper on-going performance of the ACCU-CHEK Inform II system and the competence of staff performing tests.   _________________________ <Enter the name of your facility or healthcare system> subscribes to the ____________________ <insert the name of your proficiency program.>

Survey sample handling - <insert your policy for sample handling, which should be according to the proficiency vendor’s recommendations.>

The Point of Care Coordinator delivers the samples to the nursing units where they are tested in the same manner as patients by certified operators who perform patient testing.  <Enter your practice regarding how many meters – all or a subset – are tested for each event;  if a subset is tested, include information about your practice for rotating meters through the annual cycle of survey events.>

Results are submitted to <enter the name of the proficiency vendor> without consulting with any other survey subscribers.

The survey report is reviewed by the ____<insert the title of the individual responsible for reviewing and signing off on survey reports.> Appropriate corrective action is taken in the event of survey failures.

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to conduct proficiency testing on the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| A proficiency program is selected. | <Enter the title of the healthcare professional(s) responsible for selecting and implementing a proficiency program> | <Enter the title of any policies that reference the selection of proficiency programs.> |
| Proficiency samples are sent by the proficiency vendor every _______ <enter the frequency of events>months. | <Enter the name of your Proficiency vendor> |  |
| The samples are delivered to nursing units for testing where they are tested according to routine practice on ACCU-CHEK Inform II meters by certified operators who perform patient testing. Testing is performed in the Proficiency mode of the ACCU-CHEK Inform II system. | Point-of-Care Coordinator Certified ACCU-CHEK Inform II operators who perform routine patient testing. | Performing Tests on the ACCU-CHEK Inform II System – Proficiency Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual <Enter the name of the proficiency vendor literature with testing/processing instructions.> |
| Test results are recorded on the survey results reporting form. | Point of Care Coordinator | <Enter the name of the proficiency vendor literature with testing/processing instructions> |
| The results are submitted to the proficiency vendor. | Point of Care Coordinator | <Enter the name of the proficiency vendor literature with testing/processing instructions> |
| A report is sent back from the vendor. | Proficiency vendor | <Enter the name of the proficiency vendor literature with testing/processing instructions> |
| Corrective Action is undertaken in the event of unacceptable proficiency results | <Enter the title of the healthcare professional responsible for addressing unacceptable proficiency results> | Performing Tests on the ACCU-CHEK Inform II System – Proficiency Testing Policy, Process and Procedure |

|  | Procedure |
| --- | --- |

Purpose:  The purpose of this procedure is to provide basic instructions for the system administration staff and certified operators to perform proficiency testing.  The steps in this procedure are abbreviated.  Refer to the ACCU-CHEK Inform II Operator’s Manual for more detailed testing instructions.

Acceptable Samples: <Enter the name of the proficiency vendor and the samples that are included in the survey and any special instructions included in the proficiency vendor’s instructional literature such as allowing samples to come to room temperature and mixing instructions.>

Materials:  Gather the following supplies in preparation for proficiency testing:

Note: <Insert a statement regarding where supplies are obtained to replenish stock on nursing units; example: Contact the Laboratory if you need to replenish your stock of ACCU-CHEK Inform II test strips and controls solutions or to replace a non-operational meter.>

- ACCU-CHEK Inform II meter – fully charged and coded to the test strip lot you intend to use.

- ACCU-CHEK Inform II test strip vial

- Proficiency samples

- Gloves

- Biohazard container

Procedure:

- See the proficiency vendor’s literature and the ACCU-CHEK Inform II Operator’s Manual for complete instructions on performing proficiency testing.

- <Insert your corrective action policy in the event of unacceptable proficiency results.>

Interpretation of Results:

- <Enter your policy for determining the acceptability of proficiency results.>

- Troubleshooting unacceptable results:  <Insert your policy for troubleshooting questionable results; Example text:

  - Sequester any meters demonstrating unacceptable results and a sample of the test strip lot involved.

  - Refer to the troubleshooting procedure in the “ACCU-CHEK Inform II System Maintenance” section of this manual.

  - Call ACCU-CHEK Customer Care at 1-800-440-3638 for assistance any time you have questions or concerns regarding the ACCU-CHEK Inform II system.

Review Results

|  | Policy |
| --- | --- |

The ACCU-CHEK Inform II meter system stores up to 2,000 results in the meter memory.  Results may be reviewed in the meter by any certified operator as needed in the course of patient care or system maintenance.

|  | Procedure |
| --- | --- |

- Turn on the ACCU-CHEK Inform II meter.

- Enter  your operator ID by means of <insert your policy for entering operator ID’s; e.g., barcode scanning and/or manual entry> NOTE: If the operator ID you enter is not accepted, attempt to re-enter it.  If it is still rejected, contact your supervisor or Point of Care Coordinator.  DO NOT attempt to review results under another operator’s ID.

- From the Main Menu, touch Review Results

- A chronological list of all test results are displayed on the meter screen.  The most recent results appear at the top of the list.

- Touch to display older results. Touch  to display more recent results.

- Touch any result on the list to view result details. The following information will be displayed: Patient ID, QC level or sample ID, the lot number of the reagent(s) used to perform the test, the test result, the date and time the test was performed, and comments that were entered at the time the test was performed.

- Touch  to return to the previous screen.

- Touch Patient to specify a patient whose result you wish to display.

- Barcode scan or manually enter the patient ID for the patient whose results you wish to display. Use the up and down arrow buttons to scroll through the patient’s results.

- Touch QC to review all QC results. Use the up and down arrow buttons to scroll through the QC results.

- Touch any QC result to view result details. Touch  to return to the previous screen.

- Touch  to return to the Main Menu

Transferring Data

|  | Policy |
| --- | --- |

<Enter a description of your electronic method for transferring data from the meter into the blood glucose data management system, laboratory information system, hospital information system and patient chart.  Example text follows.>

Data is transferred from the meter into <enter the name of your data management system> as follows:

- Wireless Data Transfer:  Results will automatically transfer wirelessly into <insert the name of your data management system> if a wireless connection is currently available upon completion of a test.  A wireless connection is activated when the testing operator touches   upon completing a test.  An available wireless connection is confirmed by the presence of  at the bottom of the screen.    If  is not present, a wireless connection is not currently available.  If a wireless connection is not currently available the meter will automatically send results as soon as a wireless connection is established.  The meter will connect to <enter the name of your data management system> every 10 minutes and transfer new data when a wireless connection is available even if the meter is powered off.  This connection gives the meter the most up to date settings and ADT information when available.

- Wired Data Transfer:  Results are transferred into <enter the name of your data management system> by placing the meter into a base unit that is connected to the hospital network through (select whichever is appropriate:  an RJ45 connection or a PC that is powered on and connected to the hospital network).  When properly docked the meter will automatically transfer test results.  The following series of screen displays will appear to confirm that data  transfer is occurring:

  - “Connecting,” (meter is establishing connection)

  - “Transferring,” (meter is transferring information)

  - “Idle” (meter has finished communications and is ready to use.)

The meter can be removed and used ant any time during the transfer of information.  Any information not completed will be transferred the next time the meter is docked. When properly docked the meter will communicate every 10 minutes even if the meter is off.  This connection gives the meter the most up to date settings and ADT information when available. This transfer can occur also when a wirelss meter is docked in a connected base unit.

- (This section applies to data management systems that do not have remote wireless or wired connectivity – commonly known as “sneakernet”)  Results are transferred into <enter the name of your data management system> by docking meters into a base unit that is directly connected by USB cable into a PC with the <enter the name of your data management system> running.  When properly docked the meter will automatically transfer test results.  The following series of screen displays will appear to confirm that data  transfer is occurring:

  - “Connecting,” (meter is establishing connection)

  - “Transferring,” (meter is transferring information)

  - “Busy” (communication is still active but the meter is busy processing the data received or waiting for the data message from the DMS)

  - “Idle” (meter has finished communications and is ready to use.)

The meter can be removed and used at any time during the transfer of information.  Any information not completed will be transferred the next time the meter is docked.  This connection gives the meter the most up to date settings.)

- <Enter a description of your method for transferring results into your Laboratory information system (LIS) including any queue process for holding results for review.>

- <Enter a description of your method for transferring results into patient records.>

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken in order to transfer data from the ACCU-CHEK Inform II system into various data management systems and patient records.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Procedures Followed |
| Testing is performed using ACCU-CHEK Inform II meters on nursing units. A result appears on meter screen and is stored in the meter at the end of a test. | Any operator who is currently certified to perform testing. | Performing Tests on the ACCU-CHEK Inform II System section of this manual |
| Results are uploaded from the meter into <enter the name of your data management system> by means of <enter your method of data transfer from the meter into the data management system, e.g., wireless connection, placing the meter into a base unit that is connected to the hospital network through either an RJ45 connection or a PC connected to the hospital network.> | Any operator who is currently certified to perform testing. | Performing Tests on the ACCU-CHEK Inform II System section of this manual |
| Results are transmitted from <enter the name of your glucose data management system> into <enter the name of your LIS> by means of <enter the type of interface between your glucose data management system and your LIS> | <Enter the staff position(s) responsible for establishing data transfer protocols between IT systems.> | <Enter the titles of any relevant internal IT and Laboratory protocols and procedures> |
| Results are transmitted from <enter the name of your Laboratory information system> onto patient charts by means of <enter your protocol for transmitting results from the LIS onto patient charts.> <or> Results are entered into the patient chart manually at the time of testing on nursing units. | <Enter the titles of staff positions responsible for establishing data transfer protocols between IT systems.> Any operator who is currently certified to perform testing. | <Enter the titles of any relevant internal IT, Laboratory and Medical Records protocols and procedures.> |

ACCU-CHEK Inform II System Maintenance

Storing, Maintaining and Handling ACCU-CHEK Inform II Meters and Base Units

|  | Policy and Practice |
| --- | --- |

Meters, system components and reagents are stored, maintained and handled according to manufacturer’s instructions and in compliance with all established safety and infection control policies and regulatory guidelines.

Meter and Base Unit Storage, Maintenance and Handling

  - <State your policy for storing meters on the nursing units; including whether meters should be stored in their carrying cases or in the base units when they are not being used for testing.>

  - Handle the meter and its system components carefully. Avoid dropping it or banging it.

  - Protect the base unit from dripping liquid.

  - Do not immerse the meter and base unit in any liquid.

  - Observe the following limits for temperature and humidity when storing and using the meter.

| Specification | Meter | Base Unit | Power Supply |
| --- | --- | --- | --- |
| Measurement temperature | 61–95 °F (16–35 °C) | N/A | N/A |
| Storage temperature (long-term storage) | 41 to 104 °F (5 to 40 °C) At 10-85% RH (non-condensing) | 41 to 104 °F (5 to 40 °C) At 10-85% RH (non-condensing) | 41 to 104 °F (5 to 40 °C) At 10-85% RH (non-condensing) |

  - Do not expose the meter to excessive sources of heat for prolonged periods of time when performing a test. Potential sources of heat include but are not limited to:

– Leaving the meter under a bilirubin light or photo therapy light

– Leaving the meter on a bed warmer

– Leaving the meter in an isolette

Cleaning and Disinfecting the ACCU-CHEK Inform II System Components

  - Meter cleaning and disinfecting - Cleaning and disinfecting the exterior surface of the meter is, at minimum, recommended daily for dedicated patient devices. Meters used with multiple patients may require more frequent cleaning and disinfecting. Follow recommendations from the FDA, CDC, and CMC and your facility’s policies and procedures for infection control. ,

NOTE: The FDA recommends that Point of Care testing devices, such as blood glucose meters, should be used only on one patient and not shared. If dedicating blood glucose meters to a single patient is not possible, the meters must be properly cleaned and disinfected after every use following the guidelines given in this manual.

Meters are cleaned and disinfected <insert your facility’s preference for frequency of meter cleaning and disinfecting.>

  - Base Unit cleaning and disinfecting - Base units are cleaned and disinfected <insert your facility’s preference for frequency of base unit cleaning and disinfecting.>

  - Accessory box cleaning and disinfecting - Accessory boxes are cleaned and disinfected <insert your facility’s preference for frequency of accessory box cleaning and disinfecting.>

  - Cleaning and disinfecting are companion procedures that are generally performed together at the same time.

    - Cleaning removes visible soil and organic material.

    - Disinfecting destroys most recognized pathogenic microorganisms but not necessarily all microbial forms (e.g., bacterial spores).

- NOTE:  Follow all facility safety and infection control policies when handling, cleaning and disinfecting ACCU-CHEK Inform II meters.

The product used for cleaning and disinfecting the ACCU-CHEK Inform II meters and system components is <insert your facility’s choice of product – see the ACCU-CHEK® Inform II Operator’s Manual for a list of approved materials for cleaning and disinfecting – Clorox® Germicidal Wipe EPA reg. no. 67619-12 or Super Sani-Cloth® Germicidal Disposable Wipes EPA reg. no. 9480-4))>

  - For technical assistance or questions on cleaning and disinfecting please contact ACCU-CHEK Customer Care at 1-800-440-3638.

General Cleaning and Disinfecting Procedural Considerations:

Dos:

  - Follow all facility safety and infection control policies when handling, cleaning and disinfecting ACCU-CHEK Inform II meters.

  - If you notice any signs of deterioration after cleaning and disinfecting a meter system, contact ACCU-CHEK Customer Care at 1-800-440-3638 for assistance.

  - Obtain and dispose of acceptable cleaning and disinfection products per facility guidelines.

Don’ts

  - Do Not clean or disinfect the meter while performing any type of test.

  - Do Not allow pooling of liquid on the touchscreen.

  - Do Not spray anything onto the meter or base unit.

  - Do Not immerse the meter or base unit in liquid.

  - Do Not get liquid into the strip port!  If liquid does get into the test strip port, immediately dry the components with a dry cloth or gauze pad. If solution is allowed to collect in any meter opening, severe damage to the system can occur.  If you suspect that moisture may have entered the strip, perform glucose control testing.

  - Do Not wipe the electrical connectors on the back of the base unit.

  - Do Not use any cleaning and disinfecting product other than that which is recommended by the manufacturer, identified in this procedure and provided through normal procurement policies and procedures.

Cleaning and Disinfecting the ACCU-CHEK Inform II System Components (meters, base units, accessories boxes) – Procedure

| Component | Cleaning Procedure | Disinfecting Procedure |
| --- | --- | --- |
| ACCU-CHEK Inform II Meter | - Place the meter on a level surface prior to cleaning. - Power off the meter. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it before cleaning the surfaces of the meter. - Use an approved wipe to clean by gently wiping the surfaces of the meter being careful to remove any visible soil or organic material. Use additional wipes as needed. Note: Carefully wipe around the meter test strip port area, making sure that no liquid enters the test strip port. - Dry the meter surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the meter at the completion of cleaning. | - Place the meter on a level surface prior to disinfecting. - Power off the meter. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it before disinfecting the surfaces of the meter. - Use an approved wipe to disinfect by gently wiping the surfaces of the meter three times vertically and three times horizontally. Use additional wipes as needed. Note: Carefully wipe around the meter test strip port area, making sure that no liquid enters the test strip port. - Allow the surfaces of the meter to remain damp with the disinfecting solution for <one full minute if using Clorox Germicidal Wipes or two full minutes if using Super Sani-Cloth wipes>. - Dry the meter surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the meter at the completion of cleaning. |
| ACCU-CHEK Inform II Base Unit | - Place the base unit on a level surface prior to cleaning. - Unplug the base unit. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it. - Use an approved wipe to clean by gently wiping the surfaces of the base unit being careful to remove any visible soil or organic material. Use additional wipes as needed. Note: Do not wipe the electrical connectors on the back of the base unit. - Dry the base unit surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the base unit at the completion of cleaning. | - Place the base unit on a level surface prior to disinfecting. - Unplug the base unit. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it. - Use an approved wipe to disinfect by gently wiping the surfaces of the base unit. Use additional wipes as needed. Note: Do not wipe the electrical connectors on the back of the base unit. - Allow the surfaces of the meter to remain damp with the disinfecting solution for <one full minute if using Clorox Germicidal Wipes or two full minutes if using Super Sani-Cloth wipes>. - Dry the base unit surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the base unit at the completion of cleaning. |
| ACCU-CHEK Inform II Accessory Box | - Place the accessory box on a level surface prior to cleaning. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it before cleaning the surfaces of the accessory box. - Use an approved wipe to clean by gently wiping the surfaces of the accessory box being careful to remove any visible soil or organic material. Use additional wipes as needed - Dry the accessory box surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the accessory box at the completion of cleaning. | - Place the accessory box on a level surface prior to disinfecting. - Remove an approved wipe from its packaging. Squeeze the wipe out and blot on a dry paper towel to remove any excess solution from it before disinfecting the surfaces of the accessory box. - Use an approved wipe to disinfect by gently wiping the surfaces of accessory box. Use additional wipes as needed. - Allow the surfaces of the meter to remain damp with the disinfecting solution for <one full minute if using Clorox Germicidal Wipes or two full minutes if using Super Sani-Cloth wipes>. - Dry the accessory box surfaces thoroughly with a soft cloth or gauze after cleaning. Visually verify that no solution is seen anywhere on the accessory box at the completion of cleaning. |

  - Test Strip Storage and Handling:

  - Use the test strips at temperatures between 61-95 °F (16-35 °C).

  - Use the test strips between 10-80 % relative humidity. Humidity is the amount of dampness in the air.

  - Store the test strips at temperatures between 36–86 °F (2–30 °C).Do not freeze.

  - Store unused test strips in the original container with the cap closed. Do not remove test strips from the test strip container and put them into another container such as a plastic bag or pocket, etc.

  - Close the container tightly immediately after removing a test strip to protect the test strips from humidity.

  - Use the test strip immediately after removing it from the container.

  - Discard the test strips that are past the expiration date printed on the test strip container. If the expiration date is missing or illegible, do not use the test strips.

  - Do not apply blood or control solution to the test strip before inserting it into the meter. If a result appears before applying blood or control solution, do not act on that result.

Troubleshooting the ACCU-CHEK Inform II System

<Note: This example procedure presumes that you intend to allow a basic level of troubleshooting to be performed on nursing units by testing operators.  If this is not your policy, you will need to modify this procedure extensively to reflect your policies.>

|  | Policy |
| --- | --- |

Basic Troubleshooting of ACCU-CHEK Inform II meter malfunctions or questionable performance is undertaken on nursing units by any certified operator or his/her supervisor to the extent described in the Interpretation of Results section of the Patient and Quality Control Testing Procedures in this manual.  If basic troubleshooting on the nursing unit fails to resolve the concern, the operator is to sequester the meter and test strips involved and deliver them to the <insert the name of the department responsible for advanced troubleshooting; e.g., the Laboratory> for advanced troubleshooting.

Advanced troubleshooting is performed by <insert the name of the department and staff position(s) responsible for advanced troubleshooting; e.g.; in the Laboratory by the Point-of-Care Coordinator> when basic troubleshooting steps fail to resolve an incident of meter malfunction or questionable performance.

If advanced troubleshooting fails to resolve the concern, the meter is taken out of service and replaced according to the Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure section of this manual below.

|  | Process Table |
| --- | --- |

Purpose: The following process table describes the inter-related actions that are undertaken for the purpose of resolving malfunctions or unexpected performance of the ACCU-CHEK Inform II system.

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Related Procedures |
| A policy and procedure for troubleshooting malfunctions or unexpected performance is established. | <Enter the title(s) of the healthcare professional(s) responsible for developing this manual> | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual |
| A certified operator encounters a situation in which the ACCU-CHEK Inform II system is physically broken, malfunctions or performs in an unexpected way. | The certified operator involved in the incident | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual |
| The meter and test strip vial involved are sequestered at the time of the incident. | The certified operator involved in the incident | Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure |
| The certified operator undertakes basic steps are to identify and correct the cause of the concern. If the issue is resolved by the basic troubleshooting steps taken, normal testing resumes. If the issue is not resolved, the meter and test strips involved are taken to the _______________ <enter the department you have designated to be responsible for advanced troubleshooting. Example: Laboratory) for advanced troubleshooting. | The certified operator involved in the incident <Enter the staff position(s) responsible for advanced troubleshooting; example Point-of-Care Coordinator> Call ACCU-CHEK Customer Care for assistance at 1-800-440-3638 | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure |
| The meter and test strips involved in the incident undergo advanced troubleshooting. | <Enter the staff position(s) responsible for advanced troubleshooting; example Point-of-Care Coordinator> Call ACCU-CHEK Customer Care for assistance at 1-800-440-3638 | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure |
| If the issue is resolved by advanced troubleshooting, the meter is placed back into service. If the issue is not resolved by advanced troubleshooting, the meter is replaced. | <Enter the staff position(s) responsible for comprehensive troubleshooting; example Point-of-Care Coordinator> Call ACCU-CHEK Customer Care for assistance at 1-800-440-3638 | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Elements of Quality Assurance Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure |
| Records of advanced troubleshooting including a description of the incident and the steps taken to resolve the incident are documented in the <describe the location where instrument records are kept – this could be an electronic record in the data management system or a paper record in the Elements of Quality Assurance section of this manual. Consider using the troubleshooting form below for record keeping purposes>. | <Enter the staff position(s) responsible for comprehensive troubleshooting; example Point-of-Care Coordinator> | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Elements of Quality Assurance – Policy and Documentation |

|  | Procedure |
| --- | --- |

Basic Troubleshooting Procedure

- Troubleshoot system malfunctions or questionable performance on nursing units to the extent described in the Troubleshooting section of the Patient or Quality Control Testing Procedures in this manual.

- If troubleshooting on the nursing unit fails to resolve the concern, sequester the meter and test strips involved and deliver them to the <insert the name of the department responsible for further troubleshooting> for advanced troubleshooting.

Advanced Troubleshooting Procedure

Troubleshooting Operational Issues (Power failures, operational error messages; failure to function properly) : <consider the steps listed below and insert/delete/edit/add items according to your practices and preferences.>

  - See the troubleshooting section of the ACCU-CHEK Inform II Operator’s Manual for advanced troubleshooting guidance to resolve meter operational issues.

General guidance for interpreting on-screen message and error codes:

- If the error message “Strip Defect Error” appears on the display, the test strip may be defective or the blood glucose result may be extremely low and below the meter’s measurement range.  Refer to the test strip package insert, perform a quality control test using a new test strip, review proper testing procedure, and repeat the blood glucose test, or follow your facility’s testing policy.

- If the meter displays “Type Bad Dose,” there may be insufficient amount of blood on the test strip. Repeat the test using a new test strip, ensuring proper sample application, or refer to the test strip package insert.

| All error messages displayed by the system have a letter identifying the message type, a number and a description of the error to help the operator take action to resolve the problem. The different message types are in the table below. |  |
| --- | --- |
| - E | Identifies the notification as an Error. The information notifies the operator that an error has occurred. |
| - W | Identifies the notification as a Warning. The information does not block the operator from continuing, but rather gives the operator information that may suggest an alternate workflow is required. |
| - I | Identifies the notification as Informational only. Informational notifications present the operator with contextual information, and allow the operator to proceed after confirming the notification. |
| - D | Identifies a Decision point. Decision notifications provide the operator with a choice based on contextual information. |

  - Perform a meter reset when indicated according to the troubleshooting section of the ACCU-CHEK Inform II Operator’s Manual.   See the ACCU-CHEK Inform II Operator’s Manual for instructions on performing a meter reset.

  - Record troubleshooting steps taken.

  - Call ACCU-CHEK Customer Care at 1-800-440-3638 and undertake any additional troubleshooting as directed by ACCU-CHEK Customer Care.

  - Record information collected in items 1 – 4 on a troubleshooting form <see sample form below.>

  - Keep a record of your troubleshooting in the Elements of Quality Assurance section of this manual.

Troubleshooting Questionable Results <consider the steps listed below and insert/delete/edit/add items according to your practices and preferences.>

- Begin investigation as soon as possible after the incident occurs.

- Sequester the meter and test strip vial involved.

- Review the QC and linearity history of the meter.

- Perform validation testing (e.g., precision, linearity, accuracy studies) using the meter and test strips involved in the incident.

- Interview the operator who performed the test; ask about any unusual circumstances or difficulties associated with the incident.

- Review the operator’s competency records.

- Observe the operator performing a test.

- Review the patient chart and observe the patient for signs of compromised circulation and/or environment.

- Review the patient’s lab history for any abnormal lab results – All disciplines (Chemistry, Hematology, Microbiology, Blood Bank).

- Call ACCU-CHEK Customer Care at 1-800-440-3638 and undertake any additional troubleshooting as directed by ACCU-CHEK Customer Care.

- Record information collected in items 1 – 9 on a troubleshooting form <see sample form on the next page>

- Keep a record of your troubleshooting in the Elements of Quality Assurance section of this manual.

Replacing ACCU-CHEK Inform II Meters

|  | Policy |
| --- | --- |

When basic and advanced troubleshooting fail to resolve a meter malfunction or questionable performance issue, the <insert the staff position(s) responsible for advanced troubleshooting> consults with the manufacturer to arrange replacement or repair of the meter.  <Insert a comment here regarding whether or not meters may be replaced or repaired by any other department staff.  Example text: “The replacement of all malfunctioning meters must be arranged by the Laboratory Point-of-Care Coordinator”.>

|  | Process Table |
| --- | --- |

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Related Procedures |
| Basic and Advanced troubleshooting fail to resolve a meter malfunction or questionable performance. | <Enter the staff position(s) responsible for basic and advanced troubleshooting; example Point-of-Care Coordinator> Call ACCU-CHEK Customer Care for assistance at 1-800-440-3638 | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure |
| The <enter the department and staff position(s) responsible for advanced troubleshooting> calls the manufacturer to undertake any additional troubleshooting as directed by ACCU-CHEK Customer Care. | <Enter the department and staff position(s) responsible for advanced troubleshooting.> Call ACCU-CHEK Customer Care for assistance at 1-800-440-3638 | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II test strip package insert Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Patient Testing Policy, Process and Procedure |
| Replacement of the meter is arranged. | <Insert department and staff title> ACCU-CHEK Customer Care | Troubleshooting the ACCU-CHEK Inform II System – Policy, Process and Procedure Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure |
| The replacement meter is shipped to the <insert your preferred receiving department> and then is delivered to the <enter the department and staff position(s) responsible for preparing new meters for service. Example text: The replacement meter is shipped to the Materials Management Department and is then delivered to the Laboratory Point-of-Care Coordinator> | <Insert the department(s) and staff position(s) responsible for receiving the meter, such as: Receiving, materials management, pharmacy, Laboratory, Point-of-Care Coordinator> | Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure <Enter the title of any procedures that address product receiving protocols in your facility> |
| The replacement meter undergoes assembly, charging and configuration. | Point-of-Care Coordinator | ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II System Setup/Configuration Parameters – Policy, Process and Procedure |
| The replacement meter undergoes the following validation testing <list any validation testing you choose to perform prior to placing a meter into service for patient testing; Examples: Quality control testing, linearity testing, lab analyzer correlation testing, precision testing> | Point-of-Care Coordinator | Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure ACCU-CHEK Inform II Operator’s Manual Performing Tests on the ACCU-CHEK Inform II System – Quality Control Testing Policy, Process and Procedure Performing Tests on the ACCU-CHEK Inform II System – Linearity Testing Policy, Process and Procedure ACCU-CHEK Inform II Preliminary Method Validation - General Procedure |
| Documentation of the replacement meter’s validation testing is placed in the Elements of Quality Assurance section of this manual | Point of Care Coordinator | Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure |
| The replacement meter is placed into service for patient testing, or held in reserve for service as needed. | Point-of-Care Coordinator | Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure |
| The meter that has been replaced is returned to the manufacturer as instructed by ACCU-CHEK Customer Care. | Point-of-Care Coordinator <Enter the department and staff position(s) responsible for shipping products from your facility> ACCU-CHEK Customer Care (1-800-440-3638) | Replacing ACCU-CHEK Inform II Meters – Policy, Process and Procedure <Enter any procedures you follow for shipping products from your facility> |

|  | Documentation |
| --- | --- |

<Insert copies of troubleshooting records and replacement meter validation documentation, such as Quality Control, linearity, lab correlation, precision testing, etc., in the Quality Assurance section below.>

ACCU-CHEK Inform II Elements of Quality Assurance

|  | Policy |
| --- | --- |

The following practices are undertaken at <enter the name of your facility/healthcare system> in order to monitor and ensure on-going proper performance of the ACCU-CHEK Inform II system and reducing errors throughout its usage for patient testing:

<List all Quality Assurance initiatives you would like to include below.  Consider the examples below and insert/delete/edit/add items according to your practices or preferences.>

- Preliminary method validation studies.

- In-range low and high quality control testing is required <insert your QC schedule here; example: “each day of patient testing”.>

- Quality Control results are monitored for evidence of systemic aberrations <consider what reports from your data management system you would like to include in QC monitoring such as Levy Jennings and test Strip Code Statistics).

- Linearity testing is performed <insert you established routine and schedule for linearity testing; Example text: “on every meter placed into service and on (all meters or an X% representative subset of meters) every 6 months thereafter.”>

- Proficiency Testing is performed using the <insert the name of your proficiency survey>

- New test strip lot validation <insert your policy for validating new test strip lots before placing them in service, e.g., linearity testing, lot-to-lot blood comparisons, etc.>

- Troubleshooting of incidents of meter malfunction or questionable performance is conducted according to the policies and practices in the troubleshooting section of this manual.

- Replacement meter validation studies <insert your replacement meter validation practices established in the previous section above.>

- Miscellaneous Quality Assurance Initiatives: various Quality Assurance initiatives are undertaken to address systemic concerns should they arise.

- Consider developing and including a table identifying variables that can lead to potential errors and suggested mitigating actions.  Refer to CLSI Document POCT07-A footnoted below for guidance in developing this table.

| Variable | Potential Error | Mitigating Action |
| --- | --- | --- |

|  | Documentation |
| --- | --- |

<Insert copies of Quality Assurance documentation related to items listed above at the end of this section.>

ACCU-CHEK Inform II Operator Certification/Recertification

|  | Policy |
| --- | --- |

<Insert a basic description of your operator certification/recertification policy here; Example text:  “Prior to performing any patient testing, all operators undergo comprehensive initial training format such as self paced, classroom or end unit documented observation as outlined below.  In order to ensure on-going competency, all operators then undergo recertification training as outlined below at six months following the initial certification training, at least annually thereafter and any time occasions arise indicating a need for re-training (e.g., product changes, adverse incidents involving non-compliance with proper testing practices).  Abbreviated in-servicing may occur at any time on occasions when information on minor updates needs to be disseminated to all operators.”>

All initial operator certification and recertification training is conducted by <insert the staff position(s) of qualified trainers here; e.g., a manufacturer’s representative or designee, designated individuals trained by a manufacturer’s representative or designee. <Consider keeping a list of qualified trainers in this section of your manual.>

- Initial Operator Certification Content:  <Consider the list below and include/delete/edit/add items to reflect the content of your initial certification-training program.>

- Attending an ACCU-CHEK Inform II operator certification class as outlined in the ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing.

- Certification in the following blood sample collection procedures: <enter the titles of relevant phlebotomy procedures.>

- Successful completion of a patient test under the observation of a qualified trainer

- Successful performance of low and high control tests

- Successful performance of blind sample testing

- Passing the operator certification test

- <insert any other preferred operator certification methods>

- Operator Recertification Training Content: <Consider the list below include/delete/edit/add items to reflect the content of your recertification-training program>

- Attending an ACCU-CHEK Inform II operator recertification class

- Documented review of the following blood sample collection procedures: <enter the titles of relevant phlebotomy procedures>.

- Successful completion of a patient test under the observation of a qualified trainer

- Successful performance of low and high control tests

- Successful performance of blind sample testing

- Passing a recertification test

- <insert any other preferred recertification methods>

- Operator certification/recertification training records are maintained by <insert the staff position(s) responsible for maintaining operator training records and the location where the records are stored.>

- Only a currently certified operator may perform patient testing on the ACCU-CHEK Inform II system. Operators must successfully complete the in-servicing program described in this manual.  Operators may perform tests only after entering their own ID and must under no circumstances provide their ID to enable another operator to perform testing. <enter any policy you may have regarding potential disciplinary action for testing under another operator’s ID.>

- Operators who demonstrate unwillingness or inability to utilize the ACCU-CHEK Inform II system according to proper procedures may lose their certification permanently or until they are able to demonstrate proficiency in the use of the system.

|  | Process Table |
| --- | --- |

| Process Table |  |  |
| --- | --- | --- |
| What Happens | Who is Responsible | Related Procedures |
| An operator certification program is developed for the purpose of training operators in the proper use of the ACCU-CHEK Inform II system. | <Insert the title(s) of the healthcare professional(s) responsible for developing the operator certification program; e.g., Point of Care Coordinator, Nursing educator, manufacturer’s representative.> | <Insert the title of any general facility policies and procedures regarding employee education.> ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Individuals are designated and as qualified Operator Certification/Recertification trainers | <Insert the title(s) of qualified trainers and your trainer qualifications here; e.g., a manufacturer’s representative or designee, designated individuals trained by a manufacturer’s representative or designee.> | <Insert the title of any general facility policies and procedures regarding trainer qualification.> ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Potential operators are scheduled for training sessions. <enter a description of your preferred method(s) of training; e.g., classroom training, 1:1 with trainer, on-line sessions, etc.) | <Insert the staff position(s) responsible for scheduling certification training sessions> | ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Operators attend and successfully complete a scheduled certification training session. | Individual staff position(s) whose job functions include ACCU-CHEK Inform II bedside glucose testing. Qualified staff trainers | ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Operators are added to the list of valid operators in the <insert the name of your data management system> and are assigned an operator ID and certification expiration date. | <Insert the staff position(s) responsible for maintaining operator certification records> | ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Operator certification expirations are monitored in the <insert the name of your data management system.> | <Insert the staff position(s) responsible for maintaining operator certification records> | ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Routine Recertification Process: Operators are notified _____ days <enter the number of days> prior to impending expiration by means of <insert a description of how the operators are notified; e.g.; direct communication from the Point-of-Care Coordinator or through a notification sent to the nurse manager on the operator’s nursing unit.> Non-Routine Recertification Process (product changes; correction of non-compliant practices) Operators are notified of the need for additional certification training by means of <insert a description of how the operators are notified; e.g.; direct communication from the Point-of-Care Coordinator or through a notification sent to the nurse manager on the operator’s nursing unit.> | <Insert the staff position(s) responsible for notifying operators of impending expiration.> | ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Operators participate in and successfully complete an established recertification-training program. | Individual staff position(s) whose job functions include ACCU-CHEK Inform II bedside glucose testing. Qualified staff trainers | ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Following successful recertification, each operator’s certification record is renewed in the <enter the name of your data management system> and certification expiration date is extended to reflect the next routinely required recertification training. | <Insert the staff position(s) responsible for maintaining operator certification records> | ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |
| Operators continue to re-cycle through routine and non-routine recertification training as needed for the duration of their employment in their position. | Individual staff position(s) whose job functions include ACCU-CHEK Inform II bedside glucose testing. | <Insert the title of any general facility policies and procedures regarding employee education.> ACCU-CHEK Inform II Operator’s Manual ACCU-CHEK Inform II Nursing Policies and Procedures Manual (optional) ACCU-CHEK Inform II Operator Certification/Recertification – Policy, Process and Procedure |

|  | Procedure |
| --- | --- |

- ACCU-CHEK Inform II Operator Certification/Recertification training is conducted as outlined in the following documents: <Consider the list below include/delete/edit/add items to reflect the procedural content of your certification/recertification training program.>

- ACCU-CHEK Inform II In-service Plan for Blood Glucose Testing

- <Insert the names of any other operator certification/recertification training documents/policies/procedures according to your practices.>

|  | Operator Certification/Recertification Trainers |
| --- | --- |

<Insert an on-going list of individuals who are qualified to conduct operator certification/recertification in servicing or consult IT solution.>

|  | Appendix |
| --- | --- |

<Insert miscellaneous referenced documentation here; e.g., ACCU-CHEK Inform II Operator’s Manual, ACCU-CHEK Inform II test strip and control package inserts and MSD sheets, miscellaneous manufacturer’s technical bulletins or literature.>

Notes:

Notes:

Roche Diagnostics has made all reasonable efforts to ensure that all the information contained in this manual is correct at the time of printing. However, Roche Diagnostics reserves the right to make any changes necessary without notice as part of ongoing product development.

ACCU-CHEK, ACCU-CHEK INFORM, SAFE-T-PRO are trademarks of Roche.

All other product names and trademarks are the property of their respective owners.

This document is published by Roche Diagnostics Commercial Education.

Questions or comments regarding the contents of this guide can be directed to Indianapolis.Techpub@roche.com or:

Roche Diagnostics

Commercial Education

9115 Hague Road

P.O. Box 50457

Indianapolis, IN  46250-0457

USA

This document is only available on the www.poc.roche.com and www.accu-checkinformii.com websites.

© 2012-2013, Roche Diagnostics. All rights reserved.

4302-02-0713

Revision History

Revised pages for this manual are provided by Roche Diagnostics when necessary.

| Publication Reference Number | Date | Pages Affected |
| --- | --- | --- |
| 4302-00-1012 | October 2012 | New document |
| 4302-01-0313 | March 2013 | Update critically ill statement on pages 7 and 31. |
| 4302-02-0713 | July 2013 | Update cleaning and disinfecting section to reflect addition of Super Sani-Cloth Germicidal Disposable Wipes; update method validation section on accuracy. |

CLSI Document GP2-A5; Laboratory Documents: Development and Control; Approved Guideline; Fifth Edition; 2006.

FDA Public Health Notification: Use of Fingerstick Devices on More than One Person Poses Risk for Transmitting Bloodborne Pathogens: Initial Communication, (2010).

http://www.fda.gov/ MedicalDevices/Safety/AlertsandNotices/ucm224025.htm, accessed 03/02/12.

CDC Clinical Reminder: Use of Fingerstick Devices on More than One Person Poses Risk for Transmitting Bloodborne Pathogens, (2010). http://www.cdc.gov/injectionsafety/Fingerstick-DevicesBGM.html , accessed 03/0212.

CLSI Document H4-A5 (Procedures and Devices for the Collection of Diagnostic Capillary Blood Specimens; Approved Standard – 5th Edition; 2004 was used as a resource in development of this procedure.

CLSI Document H4-A5: Procedures and Devices for the Collection of Diagnostic Blood Specimens by Skin Puncture – Approved Standard; Fifth Edition; 2004; Page 11

8 CLSI Document POCT12-A3: Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facilities; Approved Guideline – Third Edition; 2013; Section 10.2.3; Page 15.

CLSI Document H4-A5: Procedures and Devices for the Collection of Diagnostic Blood Specimens by Skin Puncture – Approved Standard; Fifth Edition; 2004; Page 11

CLSI Document POCT12-A3: Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facilities; Approved Guideline – Third Edition; 2013; Section 10.2.3; Page 15.   (nned to change to #7)

CLSI Document H4-A5: Procedures and Devices for the Collection of Diagnostic Blood Specimens by Skin Puncture – Approved Standard; Fifth Edition; 2004; Page 11

CLSI Document POCT12-A3: Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facilities; Approved Guideline – Third Edition; 2013; Section 10.2.3; Page 15.

FDA Public Health Notification: Use of Fingerstick Devices on More than One Person Poses Risk for Transmitting Bloodborne Pathogens: Initial Communication, (2010).

http://www.fda.gov/ MedicalDevices/Safety/AlertsandNotices/ucm224025.htm, accessed 03/02/12

CDC Clinical Reminder: Use of Fingerstick Devices on More than One Person Poses Risk for Transmitting Bloodborne Pathogens, (2010).

http://www.cdc.gov/injectionsafety/Fingerstick-DevicesBGM.html, accessed 03/02/12

CLSI Document POCT07-A: Quality Management: Approaches to Reducing Error at the Point of Care; Approved Guideline; October 2010.