- AmniSure® ROM Test -- Troubleshooting Checklist/Questionnaire

The checklist below is intended to help troubleshoot AmniSure ROM Test, when doubts about its results arise.  Please fax or email the completed checklist to 617- 227-2489 or info@amnisure.com . It is useful to remember the following two points:

- AmniSure is NOT a predictor of ROM. The test detects leakage of amniotic fluid into the vagina at a GIVEN POINT IN TIME.

- Discrepancies between AmniSure results and the results of older techniques of ROM diagnosis are expected. There would be no need for introducing AmniSure, a next generation reliable technology, to provide the very same results that the old unreliable methods produce.

CHECKLIST / QUESTIONNAIRE   

| Facility Name: __________________________ Contact Name:___________________________ Contact Phone #:________________________ Contact Email:___________________________ AmniSure Test Results: Positive Negative Old Method Results: Positive Negative AmniSure Lot Number: ____________ Old Methods Used: _______________________ |  |  |  |  |
| --- | --- | --- | --- | --- |
| - Possible Procedural Problems - Check: |  |  |  |  |
| Were there disinfectant solutions or medicines applied to the vagina, causing the destruction of the sample (in such a case, collecting the AmniSure swab from posterior wall or from cervix may be advisable)? | Yes |  | No |  |
| Was the swab saturated in the vagina for 1 minute? | Yes |  | No |  |
| Was the swab washed in the vial for 1 minute? | Yes |  | No |  |
| Did the swab with sample touch anything before being washed in the solvent? | Yes |  | No |  |
| Was the test strip immersed in the vial no later then 30 minutes after the sample was taken? | Yes |  | No |  |
| Was the test strip immersed in the vial until two lines were clearly visible or after 10 minutes? | Yes |  | No |  |
| Was the test read or interpreted after 15 minutes have passed since dipping the test strip into the vial? | Yes |  | No |  |
| Was there a weak test line visible (intensity of the test line should not affect the reading of the result)? | Yes |  | No |  |
| - Possible Test/Lot Problems - Check: |  |  |  |  |
| Was the test strip bent or physically damaged? | Yes |  | No |  |
| Was the control line on the test strip clearly visible? | Yes |  | No |  |
| Were the storage conditions of the test compromised? | Yes |  | No |  |
| Were Quality Checks on selected test strips performed according to the manufacturer-recommended QC manual? | Yes |  | No |  |
| Were additional AmniSure-users polled to identify similar problems, if any? If yes, please summarize the problems identified:_______________________________ ______________________________________________________________________ | Yes |  | No |  |

| - Possible Problems With Diagnosis done by Old Methods - Check: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Was patient at term? What gestational age? ________ | Yes |  | No |  |
| Was patient in labor? | Yes |  | No |  |
| Was the amount of vaginal discharge significant? | Yes |  | No |  |
| Did patient have a history of ROM? | Yes |  | No |  |
| Were urine, semen, or other contaminants present? | Yes |  | No |  |
| Was there a significant admixture of blood present? | Yes |  | No |  |
| Was any procedure that might have caused ROM performed AFTER AmniSure? | Yes |  | No |  |
| Could suspected ROM have occurred AFTER initial negative AmniSure (i.e. patient at term, ready to deliver, sudden rupture of amniotic sac)? | Yes |  | No |  |
| Were old methods repeated to confirm their true diagnosis? | Yes |  | No |  |
| Was AmniSure repeated AFTER suspected amniotic leakage/ROM occurred? | Yes |  | No |  |
| Was Indigo Carmine Amnioinfusion performed to confirm disputed ROM diagnosis? | Yes |  | No |  |
| Was additional information to confirm/reject ROM collected? | Yes |  | No |  |
| Was patient put under observation for ROM? | Yes |  | No |  |
| Was AmniSure sample collected with the aid of a speculum examination? Note: AmniSure does not require a speculum; on occasion, clinician may prefer to collect AmniSure sample from the cervix using a speculum (i.e. if disinfectants, digital exam, or medicines eliminated vaginal discharge, etc.). | Yes |  | No |  |
| Did patient go on to deliver? How many hours after AmniSure result? ________ | Yes |  | No |  |
| Was patient induced? | Yes |  | No |  |
| Was delivery vaginal? | Yes |  | No |  |

** If AmniSure test strip is inserted into the sample vial and the sample liquid does NOT seem to go up the test strip – simply flick the vial lightly with your finger and/or shake the vial a little bit to help the liquid start going up the test strip.

- COMMENTS (if none of the above-mentioned complaints apply – use this space to explain your problem):

____________________________________________________________________________________

____________________________________________________________________________________

- If you need to troubleshoot AmniSure ROM Test, you may consider the following steps:

- Download this form and save it on your desktop.

- Fill out the form.

- Send completed form back to AmniSure International LLC via fax (617-227-2489) and/or email (info@amnisure.com).  If you have any questions, please feel free to call 617-234-4441.