### POC.28 Whole Blood Glucose Using The AccuChek® Inform II

| Prepared By | Date Adopted/Replaces | Approved By |
| --- | --- | --- |
| Linda M. Green, MT | Adopted 06/30/2014 Replaces POC.13 | Edward H. Adelstein, M.D. |

| REVIEW |  |  |  |
| --- | --- | --- | --- |
| Review or Revision Date | Comments (check all that apply) |  | Approved By |
| 03/25/2015 | No changes Revision initiated – see below Available on-line |  | E.H.Adelstein, M.D. |
| 08/31/2015 | No changes Revision initiated – see below Available on-line |  | E.H.Adelstein, M.D. |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
| Date | Page # | Description of revision |  |
| 03/25/2015 | Attachment | Instructions for CPRS template |  |
| 08/31/2015 | 4 | Removed requirement to write open date on strip vials |  |

| Distribution |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Ancillary Testing Office |  |  |  | Hospital Intranet |  |  |

| Retirement Date | Approved By | Archived Location |
| --- | --- | --- |

### POC.28 Whole Blood Glucose

Using The AccuChek® Inform II

INTENDED USE

The AccuChek® Inform II is a portable instrument used to determine whole blood glucose levels.  The test system quantitatively measures glucose in venous, arterial, or fresh capillary whole blood samples from the fingertips.

The Whole Blood Glucose is a definitive test and is useful for monitoring the blood glucose levels in diabetic patients.

CAUTION:  The Food and Drug Administration (FDA) has enacted labeling changes specific to the use of multi-patient waived whole blood glucose  devices.  Under the section of “Intended Use” in the 510(k) Summary the limitations clause “The performance of this system has not been evaluated on the critically ill” has been added.  The labeling change is specifically intended to alert the end-user that these devices have limitations and are not applicable in some health care settings.  The Clinical Laboratory Improvement Amendments (CLIA) regulations specify the use of such devices on critically ill patients as “off-label” application.  At this time, the VA Central Pathology Office is awaiting further clarification of “critically ill” and testing that is in question is at the discretion of the ordering physician and to be used with caution.

It is essential that the provider consider the following when making the decision to monitor the patient glucose using the AccuChek Inform II:

  - Is the patient defined as critically ill, such as the patient is on mechanical ventilation with a mean arterial pressure less than 60 mmHg despite the use of pressors?

  - Is the patient defined as critically ill, due to decreased peripheral blood flow due to dehydration, hypotension, shock, decompensated heart failure, or peripheral arterial occlusive disease.

Some facilities may attempt to identify critically ill patients on the basis of their location, such as ICU patients, Emergency Room (ER) patients, and Operating Room (OR) patients. The VA National Enforcement Office does not recommend the identification of critically ill patients solely by location.

PRINCIPLE

The enzyme on the test strip, mutant variant of quinoprotein glucose dehydrogenase (Mut. Q-GDH), from Acinetobacter calcoaceticus, recombinant in E.coli, converts the glucose in the blood sample to gluconolactone.  This reaction creates a harmless DC electrical current that the meter interprets for the blood glucose result.  The sample and the environmental conditions are also evaluated using a small AC signal.

SPECIMEN

To perform a whole blood glucose test, a drop of fresh whole blood is necessary.  Capillary, venous, or arterial blood may be used.  The same specimen type should be used consistently throughout the patient’s care.  If a change in whole blood collection site is made, this should be  noted on the meter by the use of a comment code in order to be reported with the result.  This should be considered when reviewing the results.

Use standard biohazard precautions when collecting any blood specimen.  Wear gloves during specimen collection and while testing.  Dispose of lancets and needles in appropriate puncture resistant containers.

Use of whole blood glucose specimens should include establishment of a baseline glucose result by the clinical laboratory method and rechecked periodically thereafter.

Capillary – To obtain a sample of capillary blood, have the patient wash their hands in warm, soapy water.  Rinse well and dry completely.  If an alcohol wipe must be used, let the finger dry completely before testing.  Warming fingers can increase blood.  Using a lancet, obtain a drop of fresh capillary blood following nursing guidelines.

Arterial – To obtain a sample of arterial blood, collect the blood in a plain sterile syringe without additives or in a blood gas syringe containing lithium heparin.  Do not use a syringe with sodium heparin.  Caution should be taken to clear arterial lines before blood is drawn.  The specimen must be tested immediately to avoid glycolysis.

Venous – To obtain a sample of venous blood, collect the blood using a syringe and needle.  The specimen must be tested immediately to avoid glycolysis and clotting of the specimen

## REAGENTS – SUPPLIES – EQUIPMENT

1.	Accu-Chek Inform™ II Glucose Meter

2.	AccuChek Inform II Test Strips (supplied by the laboratory via pneumatic tube system by calling extension 53006).

  - Use test strips temperatures between 61–95 °F or (16-35 °C)  and between 10-80% relative humidity.

  - Store test strips at temperatures between 39-86 °F (4-30 °C).

  - Keep the test strips in the original container with the cap closed.  Do not remove test strips from the container and put them into another container such as a plastic bag or pocket.  Exposure to moisture may cause the monitor to reject the strip.  Moisture will cause a deterioration of the test strip reagents.

  - Reagent composition:

    - Nitrosoaniline Mediator			6.72%

    - Quinoprotein glucose dehydrogenase   15.27%

    - Pyrroloquinoline quinone 			0.14%

    - Buffer 					         34.66%

    - Stabilizer 					0.54%

    - Non-reactive ingredients		         42.66%

3.	AccuChek Inform II Control Solutions  (supplied by the laboratory via pneumatic tube system by calling extension 53006)

- Store control solutions at 39-86 °F (4-30 °C).

- The glucose control solutions are stable for three months after first opening the bottles, or until the expiration date, whichever comes first.  Date the bottles when first opened and with the open date and  new expiration date.

- Reagent composition: Glucose, Buffer, Biological Salt,  Preservative, , Nonreactive Ingredients, and FD&C Blue #1.

4.	Other supplies to perform whole blood glucose testing.

- Capillary collections – handsoap (preferred) or alcohol swabs, cotton balls or gauze, lancets-disposable single use only.  NO PENLETS!!

- Arterial – plain sterile syringes or blood gas syringe containing lithium heparin (do not use blood from syringes containing sodium heparin).

- Venous – sterile syringe and needle.

CALIBRATION (Coding)

- Each code key belongs to a single lot number and provides important information about the lot-specific properties of the test strip.  The properties of the test strips are downloaded (as a code file) from the code key using the code key reader and sent to the meter.

This procedure allows the code key information to be stored centrally in the data management system, from where it can be sent to all the meters in the facility.

- Refer to Section 6 of the Accu-Chek Inform II Operator’s Manual for further directions.

QUALITY CONTROL

- When opening a new bottle of control solution, record on the vial the date the control was opened and the date to discard the remainder of the control solution, which will be 3 months from the opened date or the printed date, whichever comes first.  Space is provided on the vials to enter this information.

Example:    Date Opened 06/29 Discard 09/29

- Two levels (Level 1 and Level 2) of quality control will be tested at the following times:

  - At a minimum of every 24 hours of Accu-Chek Inform II use. The Accu-Chek Inform II is preset to “lockout” patient testing at 0200 in all whole blood glucose testing areas.  Control tests (Level 1 and 2) must PASS before further patient testing may commence.

  - If the Accu-Chek Inform II has been dropped.

  - Immediately after a new lot number of test strips has been put into use.

  - The battery has been replaced or other maintenance has taken place.

  - When the test results contradict the patient’s clinical symptoms and the accuracy of the Accu-Chek Inform II is in doubt.

  - Any time a test strip vial has been left open.  Even short periods of exposure to humidity can affect the test strips and cause false results.

- Acceptable ranges and mean values for the glucose control solutions are entered into the AccuChek via the RALSWeb-3.  Acceptable quality control ranges will be determined statistically and may not always compare to the range listed on the test strip vial.  The operator testing the quality controls will only see a PASS or FAIL result.  All FAIL results require the entry of a comment code and further investigation.

- If the quality control results exceed the acceptable range, repeat the test after checking to insure that neither the test strips nor the control solution are outdated.  If the controls are still not within limits, use a test strip from an unopened vial and repeat the test.  Refer to the troubleshooting section for further actions to be taken and/or contact the Ancillary Testing Coordinator in Pathology for assistance.

- Each operator performing patient testing must be trained and authorized for the meter usage.  The Ancillary Testing Coordinator will maintain and provide Nursing Education with a copy of a list of all authorized operators.  Operator competency testing will be maintained on a regular basis by the Ancillary Testing Coordinator reviewing the quality control tests of each operator to verify that their results are within acceptable range.  CAP proficiency testing (performed on each meter 3 times per year) will be coordinated by the Ancillary Testing Coordinator and will be performed by randomly chosen authorized operators who normally perform patient testing.  All CAP proficiency testing records will be retained by the Ancillary Testing Coordinator for a minimum of 2 years.

- Quality control records will be kept for a minimum of two years.

- Procedure for Testing Quality Controls

- At the main menu tap on Control Test.

- Tap on the barcode icon and scan the vial of quality control solution.

- Tap the barcode icon and scan the vial of test strips.

- Remove one test strip from the vial, close the lid, and at the prompt (green downward moving arrow), place the test strip into the Test Strip Port.

- At the prompt (flashing red blood drop) introduce a drop of the control solution to the end of the test strip.

- With the resulting PASS on the screen, tap on the green check mark in the lower right corner of the screen and continue with the other level of control solution.

- Remove the test strip and discard into regular trash.

- With the FAIL result on the screen, enter a comment code as to how you intend to proceed with resolving the problem. Continue to troubleshoot until a PASS is shown upon testing.  Possible solutions in clude:

- Check to see that the yellow reagent area has air bubbles and repeat testing if they are present.

- Repeat testing using a newly opened vial of strips.

- Repeat testing with a newly opened vial of control solution.

- Contact the Ancillary Testing Coordinator for assistance.

- Enter any additional comment codes.

PATIENT TESTING

- Gather and organize the necessary equipment and supplies at the test site.

- Prior to collecting the specimen, identify the patient by the wristband ID, patient VA photo ID card, or by having the patient state their full name and SSN.  Two identifiers are to be used every time.

CAUTION:  Be sure to wear exam gloves while performing testing on the patient and to follow all precautions for bloodborne pathogens.

- The patient’s hands should be washed in warm water and soap, rinsed and dried thoroughly before testing.  This increases the circulation of the hands and fingers and will facilitate good blood flow.

- Turn on the AccuChek Inform II glucose meter.

- After the meter performs a self-check, it will proceed to the operator ID menu.  Enter the operator ID by tapping the barcode icon and laying the red laser line on the barcode area of your hospital ID badge or by manually entering your ID on the keypad, then pressing the green √..  Your ID number must be recognized as an authorized operator before any further steps may be taken.

- Tap the Patient Test  icon.

- Enter the patient ID by tapping the barcode icon and laying the red laser line on the patient ID barcode of their wristband.  The outpatient will need to have their ID manually entered on the keypad, pressing the green √.

- Scan the barcode on the vial of test strips being used.  If the code chip entry does not match the scanned entry, an error message will appear.  Refer to the trouble shooting section of the operator’s manual to resolve this problem.

- Remove a test strip from the vial and replace the cap on the vial immediately.

- When the flashing strip icon appears on the meter display, gently insert the test strip with the yellow test window facing up and the electrodes in first.

- Once the meter has detected the test strip, the flashing red drop is the prompt to apply blood to the end of the test strip.

- After ensuring the patient’s finger is clean and dry, use a new lancet to puncture the fleshy side area of the finger (See Figure A) and wipe away the first drop of blood.

Figure A.

NOTE:  Puncture only the side area of the patient fingertips.  Long term use may cause callouses and hamper the ability to use braille if needed at some point of a diabetic patient’s lifetime.

- Dose the strip at the front edge, not on top of the strip.

- Once a sufficient blood sample has been detected, the meter beeps and the measurement begins.

- When the result is displayed, a message or warning may also appear fortifying you if the result exceeds the specified limit values.  Additionally a red flashing arrow next to the test results indicates that a result is out of range.

The value is above the normal/critical upper range

The value is below the normal/critical lower range.

- Remove the test strip and dispose of it in (regular trash is allowed).  Used lancets must be discarded into a sharps container.

- Press the ON/OFF button to turn the meter off.

- Remove gloves and dispose of (regular trash is allowed).

REPORTING OF RESULTS

- A normal fasting plasma or whole blood glucose for a normal healthy patient is 72-99 mg/dL at all VISN 15 hospitals and clinics.  Diabetic patient glucose results will vary depending upon the physician’s recommended choice of therapy, i.e. diet control or insulin control.  Therefore, any action based upon the whole blood glucose results of an individual patient is solely at the discretion of the provider in charge of the patient’s diabetic or other medical care.  Upon entry of the results into the hospital computer system (CPRS), results will be reported with the normal range for healthy patients.

- Values <45 mg/dL or >500 mg/dL are considered critical.  The test must be validated by the testing personnel by repeat testing, comparison testing with a plasma glucose by the laboratory, or comparison with the patient symptoms and previously recorded results.  This validation should follow nursing protocols for patient care.

- Upon obtaining a critical result, the provider must be notified.  This notification must be documented using the CO-CRITICAL GLUCOSE POC TEST RESULT (T) template.  (See attachment 1)

- A result indicated by “RR LO” means the result may be below 30 mg/dL or a technical error may have taken place.  A result indicated by “RR HI” means the result may be above 600 mg/dL or a technical error has taken place.  When either of these codes is obtained, the test will be repeated unless clinical symptoms are comparable and the repeat test would delay emergent patient care.  Upon validation of the result, and “RR LO” or “RR HI” will be treated the same as a critical test result.

LIMITATIONS OF PROCEDURE

- The Accu-Chek Inform II test strips are for testing fresh capillary, venous, and arterial samples.  Do not use for testing body fluids of any type.

- Hematocrit should be between 10-65%.

- Lipemic samples with triglycerides in excess of 1800 mg/dL may produce elevate results.

- Blood concentrations of galactose >15 mg/dL will cause overestimation of blood glucose results.

- Intravenous administration of ascorbic acid which results in blood concentrations of ascorbic acid >3 mg/dL will cause overestimation of blood glucose results.

- If peripheral circulation is impaired, collection of capillary blood from the approved sample sites is not advised as the results might not be a true reflection of the physiological blood glucose level.  This may apply in the following circumstances:  severe dehydration as a result of diabetic ketoacidosis or due to hyperglycemic hyperosmolar non-ketotic syndrome, hypotension, shock, decompensated heart failure NYHA Class IV or peripheral arterial occlusive disease.

- The performance of this system has not been evaluated in the critically ill.  Please note cautions regarding this on page 2.

MAINTENANCE

- Please observe the following points to ensure the reliable operation of the system over the long term:

Handle the meter and its system components carefully.  Avoid dropping it or banging it.

- Protect the base unit from dripping liquid.

- Do not immerse the meter or base unit in any liquid.

- Do not expose the meter to excessive sources of heat for prolonged periods.

- Store the system and test strips in the same environment in which they are used.

- Do not store the meter in direct sunlight or under extreme temperature conditions.

- Cleaning and disinfecting the Accu-Chek Inform II system.

- The meter must be cleaned after each patient test.

- Wear gloves when cleaning the meter.

- The gloves worn during cleaning and disinfecting should be removed and hands washed thoroughly with soap and water before performing the next patient test.

- Super Sani-Cloth® Germicidal Disposable Wipes should be used for normal day-to-day cleaning after each patient test.  Squeeze excess liquid from the wipe before using.  Dry the meter thoroughly after the 2 minute disinfection time.

- Isolation patients:  Use a pre-moistened bleach disinfecting wipe provided on the cart by the patient room doorway.  Squeeze excess liquid from the wipe.  Dry the meter thoroughly after the disinfection time listed on the packet of wipes.

- Avoid getting liquid into the test strip port!

TRANSMITTING INFORMATION

- All meters should be returned to the base unit upon completion of quality control and patient testing.  This includes isolation areas and meters at bedsides in ICU.

- Place the meter back in the base unit and verify that the touchscreen moves to the Connecting, Busy, Synchronizing Database screens.  This verifies patient results are transmitting to the computer network server and information from the RALSWeb3 database is being returned to the meter.

TROUBLESHOOTING

If unable to resolve a problem with the glucose meter, please contact the Ancillary Testing Office at 53026 or 53034.  A “Loaner” meter may be obtained by taking a non-working meter to the laboratory and filling out the Accu-Chek InformII  Meter Problem form.   All areas must be filled out in order to obtain another meter.

/tmp/in/input.doc.

REFERENCES

- Accu-Chek® Inform II Operator’s Manual, 2013

- RALSWeb3, software version 4.0, Online Help Manual 2013.

- Accu-Chek® Inform II Test Strips package insert. 2013

- Accu-Chek® Inform II Control package insert. 2013

- Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facillities; Approved Guideline – Second Edition, NCCLS Document C30-A2, Vol. 22 No. 17, 2002.

- College of American Pathologists Checklists:  Point-of- Care-Testing, All Common, and Laboratory General.

- HPM 589A4-10, Bloodborne Pathogens Exposure Control Plan, March 5, 2014

- HPM 589A4-299, Personal Protective Equipment, December 11, 2012.

- HPM589A4-368, Cleaning, Disinfection, and Sterilization Guidelines, November 28, 2012.